## FTS Progress Dashboard
|  | Cryogenics Group's Progess |  |  |  |  | RTD Installation |  |  | Rack and Cabling |  | Temp. Monitor Testing
| Cryomodule SN | Coldmass Complete | Cryomodule Complete | Cryomodule Tested | Cryomodule in Tunnel | Cryomodule Installed | Received List of RTDs used | Received Post-Installation Testing Data | RTD Installation OK? | Cabling to Rack Complete | Tempature Monitor Installed | Loaded Calibration Curves | Integration Test | Notes
0.041
64117.6471 | 401 |  |  |  |  | Tunnel |  | See Notes | See Notes | LS1_N0305 |  |  |  | Installed Pairs locations have been swapped; 4 additional sensors installed
64117.6471 | 402 |  |  |  |  | Tunnel |  |  | See Notes | LS1_N0305 |  |  |  | 4 additional sensors installed
64117.6471 | 403 |  |  |  |  | Tunnel |  |  | See Notes | LS1_N0305 |  |  |  | 4 additional sensors installed
64117.6471 | 404 (spare) |  |  |  |  | Bay3
0.085 | 0.085
89000 | 801 |  |  |  |  | Tunnel |  |  | See Notes | LS1_N0602 |  |  |  | Cernox 1010 Installed
89000 | 802 |  |  |  |  | Tunnel |  |  | See Notes | LS1_N0602 |  |  |  | Cernox 1010 Installed
| 803 |  |  |  |  | Tunnel |  |  |  | LS1_N0602
| 804 |  |  |  |  | Tunnel |  |  |  | LS1_N1102
| 805 |  |  |  |  | Tunnel |  |  | See Notes | LS1_N1102 |  |  |  | A/B Swapped; one serial number is WAY off (2019 edit - the one off RTD is fine, just bad testing)
89000 | 806 |  |  |  |  | Tunnel |  |  |  | LS1_N1501
89000 | 807 |  |  |  |  | Tunnel |  |  |  | LS1_N1501
89000 | 808 |  |  |  |  | Tunnel |  |  |  | LS1_N1701
| 809 |  |  |  |  | Tunnel |  |  |  | LS1_N1701
| 810 |  |  |  |  | Tunnel |  |  |  | LS1_N2301
| 811 |  |  |  |  | Tunnel |  |  |  | LS1_N2301
| M801 |  |  |  |  | Tunnel | No RTDs Installed
M802 (spare)
Value | 0.29
60000 | 201 |  |  |  |  | Tunnel |  |  |  | LS2_N0208
60000 | 202 |  |  |  |  | Tunnel |  |  |  | LS2_N0208
60000 | 203 |  |  |  |  | Tunnel |  |  |  | LS2_N0508
60000 | 204 |  |  |  |  | Tunnel |  |  |  | LS2_N0508
60000 | 205 |  |  |  |  | Tunnel |  |  |  | LS2_N0708
60000 | 206 |  |  |  |  | Tunnel |  |  |  | LS2_N0708
60000 | 207 |  |  |  |  | Tunnel |  |  |  | LS2_N1008
60000 | 208 |  |  |  |  | Tunnel |  |  |  | LS2_N1008
| 209 |  |  |  |  | Tunnel |  |  |  | LS2_N1308
| 210 |  |  |  |  | Tunnel |  |  |  | LS2_N1308
| 211 |  |  |  |  | Tunnel |  |  |  | LS2_N1508
| 212 |  |  |  |  | Tunnel |  |  |  | LS2_N1508

## Sensor Install Data by CM
Diagnostic RTD Installation Data
Beam Line Section | CyroModule Type | CyroModule Number | RTD Location Tag | RTD Install Location | RTD Serial Number | Feedthough Port | Resistance @ room temp from datasheet | Resistance Tested after install at Port | My 2-wire test at port after installation | My 4-wire test at port after installation | Notes | Data sheet 4wire | 2Wire 4wire | Data sheet 4wire A/B Swap
LS1 | Beta .041 QWR | #1 | TX-481A | Before Solenoid 1 | X110676 | A5 | 87.0 | 103.4 | 104.36 | 85.09 |  | 1.91 | 19.27 | 0.11
LS1 | Beta .041 QWR | #1 | TX-481B | Before Solenoid 1 | X110667 | A5 | 85.2 | 105.1 | 106.47 | 86.89 |  | 1.69 | 19.58 | 0.11
LS1 | Beta .041 QWR | #1 | TX-482A | After Solenoid 1 | X110672 | A5 | 85.6 | 114.2 | 115.54 | 95.99 |  | 10.39 | 19.55 | 0.01
LS1 | Beta .041 QWR | #1 | TX-482B | After Solenoid 1 | X110673 | A5 | 96.0 | 103.7 | 104.92 | 85.55 |  | 10.45 | 19.37 | 0.05
LS1 | Beta .041 QWR | #1 | TX-483A | Before Solenoid 2 | X110669 | A5 | 86.1 | 102.4 | 103.59 | 84.07 |  | 2.03 | 19.52 | 0.03
LS1 | Beta .041 QWR | #1 | TX-483B | Before Solenoid 2 | X110671 | A5 | 84.1 | 104.5 | 105.59 | 85.98 |  | 1.88 | 19.61 | 0.12
LS1 | Beta .041 QWR | #1 | TX-484A | After Solenoid 2 | X110668 | A5 | 88.4 | 105.9 | 106.98 | 87.21 |  | 1.19 | 19.77 | 0.09
LS1 | Beta .041 QWR | #1 | TX-484B | After Solenoid 2 | X110670 | A5 | 87.3 | 106.9 | 108.1 | 88.33 |  | 1.03 | 19.77 | 0.07
LS1 | Beta .041 QWR | #1 | TX-485A | After Solenoid 3 | X110635 | A6 | 85.6 | 106.7 | 107.88 | 88.02 |  | 2.42 | 19.86 | 0.08
LS1 | Beta .041 QWR | #1 | TX-485B | After Solenoid 3 | X110678 | A6 | 88.1 | 104.3 | 105.52 | 85.47 |  | 2.63 | 20.05 | 0.13
LS1 | Beta .041 QWR | #1 | TX-486A | After Solenoid 3 | X110648 | A6 | 92.9 | 104.5 | 105.59 | 85.55 |  | 7.35 | 20.04 | 0.05
LS1 | Beta .041 QWR | #1 | TX-486B | After Solenoid 3 | X110677 | A6 | 85.6 | 111.6 | 112.89 | 92.8 |  | 7.2 | 20.09 | 0.1
LS1 | Beta .041 QWR | #2 | TX-481A | Before Solenoid 1 | X112419 | A5 | 77.4 | 95.8 | 96.91 | 77.43 |  | 0.03 | 19.48 | 10.43
LS1 | Beta .041 QWR | #2 | TX-481B | Before Solenoid 1 | X112432 | A5 | 67 | 85.4 | 86.46 | 66.99 |  | 0.01 | 19.47 | 10.41
LS1 | Beta .041 QWR | #2 | TX-482A | After Solenoid 1 | X112434 | A5 | 66.4 | 85.3 | 86.21 | 66.37 |  | 0.03 | 19.84 | 1.77
LS1 | Beta .041 QWR | #2 | TX-482B | After Solenoid 1 | X112373 | A5 | 64.6 | 83.6 | 84.54 | 64.72 |  | 0.12 | 19.82 | 1.68
LS1 | Beta .041 QWR | #2 | TX-483A | Before Solenoid 2 | X112430 | A5 | 79.8 | 98.1 | 99.21 | 79.78 |  | 0.02 | 19.43 | 13.28
LS1 | Beta .041 QWR | #2 | TX-483B | Before Solenoid 2 | X112379 | A5 | 66.5 | 84.9 | 85.89 | 66.4 |  | 0.1 | 19.49 | 13.4
LS1 | Beta .041 QWR | #2 | TX-484A | After Solenoid 2 | X112377 | A5 | 76.6 | 95.2 | 96.38 | 76.66 |  | 0.06 | 19.72 | 9.96
LS1 | Beta .041 QWR | #2 | TX-484B | After Solenoid 2 | X112431 | A5 | 66.7 | 85.5 | 86.53 | 66.76 |  | 0.06 | 19.77 | 9.84
LS1 | Beta .041 QWR | #2 | TX-485A | After Solenoid 3 | X112389 | A6 | 75.7 | 94.7 | 95.72 | 75.68 |  | 0.02 | 20.04 | 0.48
LS1 | Beta .041 QWR | #2 | TX-485B | After Solenoid 3 | X112392 | A6 | 75.2 | 94.4 | 95.43 | 75.23 |  | 0.03 | 20.2 | 0.47
LS1 | Beta .041 QWR | #2 | TX-486A | After Solenoid 3 | X112388 | A6 | 73.9 | 93.4 | 94.42 | 73.87 |  | 0.03 | 20.55 | 5.43
LS1 | Beta .041 QWR | #2 | TX-486B | After Solenoid 3 | X112427 | A6 | 79.3 | 98.3 | 99.43 | 79.35 |  | 0.05 | 20.08 | 5.45
LS1 | Beta .041 QWR | #3 | TX-481A | Before Solenoid 1 | x112408 | A5 | 75.2 | 94.5 | 95.2 | 75.17 |  | 0.03 | 20.03 | 6.83
LS1 | Beta .041 QWR | #3 | TX-481B | Before Solenoid 1 | x112391 | A5 | 82 | 100.8 | 101.6 | 82 |  | 0 | 19.6 | 6.8
LS1 | Beta .041 QWR | #3 | TX-482A | After Solenoid 1 | x112375 | A5 | 76.6 | 95.7 | 96.51 | 76.54 |  | 0.06 | 19.97 | 2.04
LS1 | Beta .041 QWR | #3 | TX-482B | After Solenoid 1 | x112390 | A5 | 74.5 | 93.2 | 94.09 | 74.47 |  | 0.03 | 19.62 | 2.13
LS1 | Beta .041 QWR | #3 | TX-483A | Before Solenoid 2 | x112409 | A5 | 78.2 | 96.9 | 97.64 | 78.18 |  | 0.02 | 19.46 | 5.68
LS1 | Beta .041 QWR | #3 | TX-483B | Before Solenoid 2 | x112406 | A5 | 72.5 | 91.2 | 91.96 | 72.45 |  | 0.05 | 19.51 | 5.75
LS1 | Beta .041 QWR | #3 | TX-484A | After Solenoid 2 | x112417 | A5 | 81 | 100.3 | 100.99 | 80.99 |  | 0.01 | 20 | 5.89
LS1 | Beta .041 QWR | #3 | TX-484B | After Solenoid 2 | x112393 | A5 | 75.1 | 94.5 | 95.17 | 75.07 |  | 0.03 | 20.1 | 5.93
LS1 | Beta .041 QWR | #3 | TX-485A | After Solenoid 3 | x112412 | A6 | 74.8 | 94.5 | 94.14 | 74.77 |  | 0.03 | 19.37 | 5.23
LS1 | Beta .041 QWR | #3 | TX-485B | After Solenoid 3 | x112428 | A6 | 80 | 99.5 | 100.26 | 80.02 |  | 0.02 | 20.24 | 5.22
LS1 | Beta .041 QWR | #3 | TX-486A | After Solenoid 3 | x112413 | A6 | 75.9 | 95.1 | 96.06 | 75.9 |  | 0 | 20.16 | 0.1
LS1 | Beta .041 QWR | #3 | TX-486B | After Solenoid 3 | x112429 | A6 | 75.8 | 94.8 | 95.68 | 75.82 |  | 0.02 | 19.86 | 0.08
LS1 | Beta .085 QWR | #1 | TX-881A | Before Solenoid 1 | X105264 | A5 | 34.872 | 54.7 | 55.55 | 34.84 | Cernox-1010 | 0.032 | 20.71 | 0.718
LS1 | Beta .085 QWR | #1 | TX-881B | Before Solenoid 1 | X105214 | A5 | 34.122 | 53.7 | 54.56 | 34.09 | Cernox-1010 | 0.032 | 20.47 | 0.782
LS1 | Beta .085 QWR | #1 | TX-882A | After Solenoid 1 | X105288 | A5 | 35.76 | 56 | 56.74 | 35.72 | Cernox-1010 | 0.04 | 21.02 | 0.23
LS1 | Beta .085 QWR | #1 | TX-882B | After Solenoid 1 | X105252 | A5 | 35.49 | 55.6 | 56.44 | 35.45 | Cernox-1010 | 0.04 | 20.99 | 0.31
LS1 | Beta .085 QWR | #1 | TX-883A | Before Solenoid 2 | X105229 | A5 | 34.14 | 53 | 19.98 | 0.0008 | Cernox-1010; SHORTED RTD | 34.1392 | 19.9792 | 35.3722
LS1 | Beta .085 QWR | #1 | TX-883B | Before Solenoid 2 | X105211 | A5 | 35.373 | 54.5 | 55.35 | 35.35 | Cernox-1010 | 0.023 | 20 | 1.21
LS1 | Beta .085 QWR | #1 | TX-884A | After Solenoid 2 | X105275 | A5 | 35.697 | 54.9 | 55.75 | 35.68 | Cernox-1010 | 0.017 | 20.07 | 1.12
LS1 | Beta .085 QWR | #1 | TX-884B | After Solenoid 2 | X105251 | A5 | 34.56 | 53.3 | 54.54 | 34.54 | Cernox-1010 | 0.02 | 20 | 1.157
LS1 | Beta .085 QWR | #1 | TX-885A | Before Solenoid 3 | X105286 | A6 | 34.325 | 54.1 | 54.98 | 34.3 | Cernox-1010 | 0.025 | 20.68 | 1.441
LS1 | Beta .085 QWR | #1 | TX-885B | Before Solenoid3 | X105287 | A6 | 35.741 | 55.2 | 56.4 | 34.71 | Cernox-1010 | 1.031 | 21.69 | 0.385
LS1 | Beta .085 QWR | #1 | TX-886A | After Solenoid 3 | X105204 | A6 | 34.881 | 54.7 | 55.79 | 34.82 | Cernox-1010 | 0.061 | 20.97 | 0.318
LS1 | Beta .085 QWR | #1 | TX-886B | After Solenoid 3 | X105218 | A6 | 35.138 | 53.9 | 55.05 | 35.12 | Cernox-1010 | 0.018 | 19.93 | 0.239
LS1 | Beta .085 QWR | #2 | TX-881A | Before Solenoid 1 | x110458 | A5 | 29.884 | 40.8 | 50.376 | 30.458 |  | 0.574 | 19.918 | 0.032
LS1 | Beta .085 QWR | #2 | TX-881B | Before Solenoid 1 | x110529 | A5 | 30.426 | 40.2 | 49.851 | 29.912 |  | 0.514 | 19.939 | 0.028
LS1 | Beta .085 QWR | #2 | TX-882A | After Solenoid 1 | x110469 | A5 | 30.679 | 41.7 | 51.202 | 31.443 |  | 0.764 | 19.759 | 0.03
LS1 | Beta .085 QWR | #2 | TX-882B | After Solenoid 1 | x110452 | A5 | 31.413 | 40 | 50.65 | 30.711 |  | 0.702 | 19.939 | 0.032
LS1 | Beta .085 QWR | #2 | TX-883A | Before Solenoid 2 | x110574 | A5 | 30.355 | 40.8 | 50.382 | 30.846 |  | 0.491 | 19.536 | 0.03
LS1 | Beta .085 QWR | #2 | TX-883B | Before Solenoid 2 | x110460 | A5 | 30.816 | 40.3 | 49.852 | 30.379 |  | 0.437 | 19.473 | 0.024
LS1 | Beta .085 QWR | #2 | TX-884A | After Solenoid 2 | x110463 | A5 | 30.689 | 44.7 | 54.754 | 34.491 |  | 3.802 | 20.263 | 0.022
LS1 | Beta .085 QWR | #2 | TX-884B | After Solenoid 2 | x105224 | A5 | 34.469 | 40.7 | 50.333 | 30.718 |  | 3.751 | 19.615 | 0.029
LS1 | Beta .085 QWR | #2 | TX-885A | Before Solenoid 3 | x110545 | A6 | 31.017 | 40.8 | 50.451 | 30.515 |  | 0.502 | 19.936 | 0.032
LS1 | Beta .085 QWR | #2 | TX-885B | Before Solenoid3 | x110575 | A6 | 30.483 | 41.5 | 51.253 | 31.041 |  | 0.558 | 20.212 | 0.024
LS1 | Beta .085 QWR | #2 | TX-886A | After Solenoid 3 | x109869 | A6 | 31.434 | 40.6 | 50.766 | 30.178 |  | 1.256 | 20.588 | 0.027
LS1 | Beta .085 QWR | #2 | TX-886B | After Solenoid 3 | x110456 | A6 | 30.151 | 41.9 | 51.986 | 31.459 |  | 1.308 | 20.527 | 0.025
LS1 | Beta .085 QWR | #3 | TX-881A | Before Solenoid 1 | x110650 | A5 | 85.6 | 94.1 |  |  |  | 85.6 | 0 | 63.5
LS1 | Beta .085 QWR | #3 | TX-881B | Before Solenoid 1 | x114273 | A5 | 63.5 | 72.2 |  |  |  | 63.5 | 0 | 85.6
LS1 | Beta .085 QWR | #3 | TX-882A | After Solenoid 1 | x112401 | A5 | 66.8 | 75.5 |  |  |  | 66.8 | 0 | 63.1
LS1 | Beta .085 QWR | #3 | TX-882B | After Solenoid 1 | x114274 | A5 | 63.1 | 71.7 |  |  |  | 63.1 | 0 | 66.8
LS1 | Beta .085 QWR | #3 | TX-883A | Before Solenoid 2 | x110651 | A5 | 85.5 | 94.2 |  |  |  | 85.5 | 0 | 53.5
LS1 | Beta .085 QWR | #3 | TX-883B | Before Solenoid 2 | x113702 | A5 | 53.5 | 62.4 |  |  |  | 53.5 | 0 | 85.5
LS1 | Beta .085 QWR | #3 | TX-884A | After Solenoid 2 | x110664 | A5 | 83.8 | 92.6 |  |  |  | 83.8 | 0 | 85.3
LS1 | Beta .085 QWR | #3 | TX-884B | After Solenoid 2 | x110653 | A5 | 85.3 | 93.9 |  |  |  | 85.3 | 0 | 83.8
LS1 | Beta .085 QWR | #3 | TX-885A | Before Solenoid 3 | x110665 | A6 | 87.8 | 96.5 | 107.726 | 87.857 |  | 0.057 | 19.869 | 5.143
LS1 | Beta .085 QWR | #3 | TX-885B | Before Solenoid3 | x110652 | A6 | 93 | 101.6 | 112.82 | 93.033 |  | 0.033 | 19.787 | 5.233
LS1 | Beta .085 QWR | #3 | TX-886A | After Solenoid 3 | x112416 | A6 | 76.2 | 94.9 | 96.669 | 76.332 | Remeasured Post; now OK | 0.132 | 20.337 | 9.668
LS1 | Beta .085 QWR | #3 | TX-886B | After Solenoid 3 | x110666 | A6 | 86 | 94.8 | 106.511 | 86.066 |  | 0.066 | 20.445 | 9.866
LS1 | Beta .085 QWR | #4 | TX-881A | Before Solenoid 1 | x113479 | A5 | 56.3 | 77.4 | 77.06 | 56.26 |  | 0.04 | 20.8 | 6.34
LS1 | Beta .085 QWR | #4 | TX-881B | Before Solenoid 1 | x113378 | A5 | 62.6 | 83.4 | 83.14 | 62.61 |  | 0.01 | 20.53 | 6.31
LS1 | Beta .085 QWR | #4 | TX-882A | After Solenoid 1 | x113485 | A5 | 54 | 77.4 | 74.17 | 53.93 |  | 0.07 | 20.24 | 6.57
LS1 | Beta .085 QWR | #4 | TX-882B | After Solenoid 1 | x113382 | A5 | 60.5 | 81.2 | 80.97 | 60.81 |  | 0.31 | 20.16 | 6.81
LS1 | Beta .085 QWR | #4 | TX-883A | Before Solenoid 2 | x113446 | A5 | 67.4 | 84.3 | 82.94 | 67.28 |  | 0.12 | 15.66 | 12.08
LS1 | Beta .085 QWR | #4 | TX-883B | Before Solenoid 2 | x113481 | A5 | 55.2 | 74.7 | 74.45 | 55.16 |  | 0.04 | 19.29 | 12.24
LS1 | Beta .085 QWR | #4 | TX-884A | After Solenoid 2 | x113381 | A5 | 62.2 | 82.4 | 82.17 | 62.2 |  | 0 | 19.97 | 1.4
LS1 | Beta .085 QWR | #4 | TX-884B | After Solenoid 2 | x113474 | A5 | 60.8 | 80.9 | 80.65 | 60.75 |  | 0.05 | 19.9 | 1.45
LS1 | Beta .085 QWR | #4 | TX-885A | Before Solenoid 3 | x113511 | A6 | 61.9 | 82.6 | 82.29 | 61.93 |  | 0.03 | 20.36 | 1.87
LS1 | Beta .085 QWR | #4 | TX-885B | Before Solenoid3 | x113436 | A6 | 63.8 | 84.3 | 83.91 | 63.72 |  | 0.08 | 20.19 | 1.82
LS1 | Beta .085 QWR | #4 | TX-886A | After Solenoid 3 | x113377 | A6 | 61.2 | 82 | 81.9 | 61.21 |  | 0.01 | 20.69 | 0.21
LS1 | Beta .085 QWR | #4 | TX-886B | After Solenoid 3 | x113512 | A6 | 61 | 81.8 | 81.75 | 61.01 |  | 0.01 | 20.74 | 0.19
LS2 | Beta .29 HWR | #1 | TX-281A | Before Solenoid 1 | x113471 |  | 62.1 | 80 | 82.149 | 62.074 |  | 0.026 | 20.075 | 12.826
LS2 | Beta .29 HWR | #1 | TX-281B | Before Solenoid 1 | x112355 |  | 74.9 | 92.3 | 94.441 | 74.789 |  | 0.111 | 19.652 | 12.689
LS2 | Beta .29 HWR | #1 | TX-282A | After Solenoid 1 | x113404 |  | 62.2 | 80.1 | 82.088 | 62.227 |  | 0.027 | 19.861 | 0.727
LS2 | Beta .29 HWR | #1 | TX-282B | After Solenoid 1 | x113469 |  | 61.5 | 79.6 | 81.758 | 61.551 |  | 0.051 | 20.207 | 0.649
LS1 | Beta .085 QWR | #5 | TX-881A | Before Solenoid 1 | x113441 | A5 | 69.16 | 82.2 Ω, ok | 81.97 | 61.41 |  | 7.75 | 20.56 | 0.02
LS1 | Beta .085 QWR | #5 | TX-881B | Before Solenoid 1 | x113477 | A5 | 61.39 | 90.1 Ω, ok | 89.87 | 69.11 |  | 7.72 | 20.76 | 0.05
LS1 | Beta .085 QWR | #5 | TX-882A | After Solenoid 1 | x112420 | A5 | 70.44 | 90.7 Ω, ok | 90.48 | 70.02 |  | 0.42 | 20.46 | 9.75
LS1 | Beta .085 QWR | #5 | TX-882B | After Solenoid 1 | x113380 | A5 | 60.27 | 90.6 Ω, ok | 90.36 | 70.33 |  | 10.06 | 20.03 | 0.11
LS1 | Beta .085 QWR | #5 | TX-883A | Before Solenoid 2 | x113510 | A5 | 60.21 | 84.7 Ω, ok | 84.49 | 64.2 |  | 3.99 | 20.29 | 0.08
LS1 | Beta .085 QWR | #5 | TX-883B | Before Solenoid 2 | x113440 | A5 | 64.28 | 80.5 Ω, ok | 80.31 | 60.25 |  | 4.03 | 20.06 | 0.04
LS1 | Beta .085 QWR | #5 | TX-884A | After Solenoid 2 | x113442 | A5 | 65.55 | 76.7 Ω, ok | 76.4 | 56.43 |  | 9.12 | 19.97 | 0.02
LS1 | Beta .085 QWR | #5 | TX-884B | After Solenoid 2 | x113494 | A5 | 56.45 | 85.9 Ω, ok | 85.77 | 65.48 |  | 9.03 | 20.29 | 0.07
LS1 | Beta .085 QWR | #5 | TX-885A | Before Solenoid 3 | x113492 | A6 | 55.07 | 75.6 Ω, ok | 75.37 | 55.01 |  | 0.06 | 20.36 | 0.02
LS1 | Beta .085 QWR | #5 | TX-885B | Before Solenoid3 | x113498 | A6 | 54.99 | 75.7 Ω, ok | 75.44 | 55.08 |  | 0.09 | 20.36 | 0.01
LS1 | Beta .085 QWR | #5 | TX-886A | After Solenoid 3 | x113496 | A6 | 53.98 | 82.7 Ω, ok | 82.5 | 61.67 |  | 7.69 | 20.83 | 0.05
LS1 | Beta .085 QWR | #5 | TX-886B | After Solenoid 3 | x113475 | A6 | 61.62 | 74.8 Ω, ok | 74.55 | 53.97 |  | 7.65 | 20.58 | 0.01
LS1 | Beta .085 QWR | #6 | TX-881A | Before Solenoid 1 | x112348 | A5 | 76.7115 | 97.4 Ω, OK | 97.21 | 76.57 |  | 0.14148 | 20.64 | 2.0762
LS1 | Beta .085 QWR | #6 | TX-881B | Before Solenoid 1 | x112347 | A5 | 74.4938 | 95.1 Ω, OK | 95.13 | 74.36 |  | 0.13381 | 20.77 | 2.3515
LS1 | Beta .085 QWR | #6 | TX-882A | After Solenoid 1 | x112346 | A5 | 70.711 | 91.1 Ω, OK | 90.94 | 70.58 |  | 0.13096 | 20.36 | 2.9845
LS1 | Beta .085 QWR | #6 | TX-882B | After Solenoid 1 | x112276 | A5 | 73.5645 | 94.0 Ω, OK | 93.87 | 73.44 |  | 0.1245 | 20.43 | 2.729
LS1 | Beta .085 QWR | #6 | TX-883A | Before Solenoid 2 | x112350 | A5 | 68.9416 | 88.6 Ω, OK | 88.51 | 68.82 |  | 0.12161 | 19.69 | 17.0082
LS1 | Beta .085 QWR | #6 | TX-883B | Before Solenoid 2 | x110679 | A5 | 85.8282 | 105.6 Ω, OK | 105.51 | 85.66 |  | 0.16824 | 19.85 | 16.7184
LS1 | Beta .085 QWR | #6 | TX-884A | After Solenoid 2 | x112343 | A5 | 77.7049 | 98.0 Ω, OK | 97.59 | 77.57 |  | 0.1349 | 20.02 | 0.79303
LS1 | Beta .085 QWR | #6 | TX-884B | After Solenoid 2 | x112351 | A5 | 78.363 | 98.2 Ω, OK | 97.97 | 78.22 |  | 0.14303 | 19.75 | 0.5151
LS1 | Beta .085 QWR | #6 | TX-885A | Before Solenoid 3 | x112349 | A6 | 74.5603 | 95.0 Ω, OK | 94.67 | 74.43 |  | 0.13035 | 20.24 | 4.3903
LS1 | Beta .085 QWR | #6 | TX-885B | Before Solenoid3 | x112345 | A6 | 78.8203 | 99.4 Ω, OK | 99.12 | 78.68 |  | 0.14033 | 20.44 | 4.1197
LS1 | Beta .085 QWR | #6 | TX-886A | After Solenoid 3 | x112422 | A6 | 74.7496 | 95.3 Ω, OK | 95.16 | 74.62 |  | 0.12959 | 20.54 | 4.4
LS1 | Beta .085 QWR | #6 | TX-886B | After Solenoid 3 | x112421 | A6 | 79.02 | 99.0 Ω, OK | 99.73 | 78.89 |  | 0.12997 | 20.84 | 4.1404
LS1 | Beta .085 QWR | #7 | TX-881A | Before Solenoid 1 | x112425 | A5 | 74.9701 | 96.0 Ω, OK | 95.84 | 74.89 |  | 0.08011 | 20.95 | 0.71777
LS1 | Beta .085 QWR | #7 | TX-881B | Before Solenoid 1 | x112426 | A5 | 74.1722 | 95.3 Ω, OK | 95.08 | 74.09 |  | 0.08223 | 20.99 | 0.88011
LS1 | Beta .085 QWR | #7 | TX-882A | After Solenoid 1 | x112424 | A5 | 70.5614 | 90.5 Ω, OK | 90.52 | 70.48 |  | 0.08145 | 20.04 | 6.399
LS1 | Beta .085 QWR | #7 | TX-882B | After Solenoid 1 | x113445 | A5 | 64.081 | 84.7 Ω, OK | 84.47 | 64.02 |  | 0.06098 | 20.45 | 6.5415
LS1 | Beta .085 QWR | #7 | TX-883A | Before Solenoid 2 | x113513 | A5 | 62.7742 | 82.8 Ω, OK | 82.28 | 62.83 |  | 0.05584 | 19.45 | 9.8718
LS1 | Beta .085 QWR | #7 | TX-883B | Before Solenoid 2 | x113482 | A5 | 52.9582 | 73.5 Ω, OK | 73.23 | 52.92 |  | 0.03816 | 20.31 | 9.8542
LS1 | Beta .085 QWR | #7 | TX-884A | After Solenoid 2 | x113379 | A5 | 61.2654 | 80.6 Ω, OK | 80.58 | 61.28 |  | 0.01457 | 19.3 | 0.39871
LS1 | Beta .085 QWR | #7 | TX-884B | After Solenoid 2 | x113470 | A5 | 61.6787 | 81.9 Ω, OK | 81.86 | 61.71 |  | 0.03129 | 20.15 | 0.44457
LS1 | Beta .085 QWR | #7 | TX-885A | Before Solenoid 3 | x113489 | A6 | 55.281 | 76.1 Ω, OK | 75.95 | 55.26 |  | 0.02101 | 20.69 | 9.9949
LS1 | Beta .085 QWR | #7 | TX-885B | Before Solenoid3 | x113452 | A6 | 65.2549 | 85.5 Ω, OK | 85.52 | 65.2 |  | 0.05495 | 20.32 | 9.919
LS1 | Beta .085 QWR | #7 | TX-886A | After Solenoid 3 | x113490 | A6 | 55.0027 | 76.1 Ω, OK | 76.05 | 54.98 |  | 0.02273 | 21.07 | 4.9505
LS1 | Beta .085 QWR | #7 | TX-886B | After Solenoid 3 | x113476 | A6 | 59.9305 | 80.0 Ω, OK | 80.91 | 59.93 |  | 0.00048 | 20.98 | 4.9273
LS1 | Beta .085 QWR | #8 | TX-881A | Before Solenoid 1 | x113502 | A5 | 62.5383 | 83.2 Ω, OK | 83.05 | 62.24 |  | 0.29834 | 20.81 | 0.38
LS1 | Beta .085 QWR | #8 | TX-881B | Before Solenoid 1 | x113405 | A5 | 62.62 | 83.4 Ω, OK | 83.08 | 62.29 |  | 0.33 | 20.79 | 0.24834
LS1 | Beta .085 QWR | #8 | TX-882A | After Solenoid 1 | x113483 | A5 | 53.3084 | 73.3 Ω, OK | 73.15 | 52.7 |  | 0.60837 | 20.45 | 9.8
LS1 | Beta .085 QWR | #8 | TX-882B | After Solenoid 1 | x113472 | A5 | 62.5 | 82.7 Ω, OK | 82.37 | 62.11 |  | 0.39 | 20.26 | 8.8016
LS1 | Beta .085 QWR | #8 | TX-883A | Before Solenoid 2 | x113450 | A5 | 64.1492 | 84.2 Ω, OK | 84 | 63.75 |  | 0.39923 | 20.25 | 4.3603
LS1 | Beta .085 QWR | #8 | TX-883B | Before Solenoid 2 | x113443 | A5 | 68.1103 | 87.9 Ω, OK | 87.61 | 67.67 |  | 0.44033 | 19.94 | 3.5208
LS1 | Beta .085 QWR | #8 | TX-884A | After Solenoid 2 | x113454 | A5 | 63.8578 | 83.4 Ω, OK | 83.22 | 63.11 |  | 0.7478 | 20.11 | 0.70065
LS1 | Beta .085 QWR | #8 | TX-884B | After Solenoid 2 | x113473 | A5 | 62.4094 | 82.2 Ω, OK | 82.04 | 62.06 |  | 0.34935 | 19.98 | 1.7978
LS1 | Beta .085 QWR | #8 | TX-885A | Before Solenoid 3 | x113448 | A6 | 66.0212 | 85.8 Ω, OK | 85.81 | 65.61 |  | 0.41123 | 20.2 | 0.72339
LS1 | Beta .085 QWR | #8 | TX-885B | Before Solenoid3 | x112415 | A6 | 66.3334 | 85.9 Ω, OK | 85.85 | 65.9 |  | 0.43339 | 19.95 | 0.12123
LS1 | Beta .085 QWR | #8 | TX-886A | After Solenoid 3 | x112356 | A6 | 78.1917 | 98.3 Ω, OK | 98.29 | 77.57 |  | 0.62175 | 20.72 | 3.4173
LS1 | Beta .085 QWR | #8 | TX-886B | After Solenoid 3 | x112403 | A6 | 80.9873 | 101.3 Ω, OK | 101.37 | 80.36 |  | 0.62732 | 21.01 | 2.1683
LS1 | Beta .085 QWR | #9 | TX-881A | Before Solenoid 1 | x113488 | A5 | 55.2895 | 55.4 Ω, OK | 75.86 | 55.15 |  | 0.13953 | 20.71 | 5.3936
LS1 | Beta .085 QWR | #9 | TX-881B | Before Solenoid 1 | x113376 | A5 | 60.5436 | 72.7 Ω, OK | 81 | 60.4 |  | 0.14364 | 20.6 | 5.1105
LS1 | Beta .085 QWR | #9 | TX-882A | After Solenoid 1 | x113478 | A5 | 54.4246 | 54.6 Ω, OK | 74.69 | 54.25 |  | 0.17465 | 20.44 | 9.0543
LS1 | Beta .085 QWR | #9 | TX-882B | After Solenoid 1 | x113444 | A5 | 63.3043 | 63.5 Ω, OK | 83.41 | 63.07 |  | 0.23427 | 20.34 | 8.6454
LS1 | Beta .085 QWR | #9 | TX-883A | Before Solenoid 2 | x113484 | A5 | 55.4457 | 55.4 Ω, OK | 75.26 | 55.26 |  | 0.18569 | 20 | 7.5631
LS1 | Beta .085 QWR | #9 | TX-883B | Before Solenoid 2 | x113383 | A5 | 62.8231 | 62.9 Ω, OK | 82.38 | 62.66 |  | 0.16308 | 19.72 | 7.2143
LS1 | Beta .085 QWR | #9 | TX-884A | After Solenoid 2 | x113491 | A5 | 54.856 | 54.8 Ω, OK | 74.52 | 54.65 |  | 0.20596 | 19.87 | 11.2308
LS1 | Beta .085 QWR | #9 | TX-884B | After Solenoid 2 | x113449 | A5 | 65.8808 | 65.7 Ω, OK | 85.74 | 65.6 |  | 0.28077 | 20.14 | 10.744
LS1 | Beta .085 QWR | #9 | TX-885A | Before Solenoid 3 | x113486 | A6 | 52.7927 | 52.9 Ω, OK | 73.11 | 52.57 |  | 0.22275 | 20.54 | 8.7695
LS1 | Beta .085 QWR | #9 | TX-885B | Before Solenoid3 | x113514 | A6 | 61.3395 | 73.7 Ω, OK | 81.59 | 61.15 |  | 0.18953 | 20.44 | 8.3573
LS1 | Beta .085 QWR | #9 | TX-886A | After Solenoid 3 | x113468 | A6 | 62.7578 | 63.0 Ω, OK | 83.68 | 62.55 |  | 0.20785 | 21.13 | 0.58081
LS1 | Beta .085 QWR | #9 | TX-886B | After Solenoid 3 | x113384 | A6 | 61.9692 | 62.3 Ω, OK | 82.58 | 61.75 |  | 0.21919 | 20.83 | 1.0078
LS1 | Beta .085 QWR | #10 | TX-881A | Before Solenoid 1 | x111769 | A5 | 84.6108 | 99.6Ω, OK | 105.68 | 84.84 |  | 0.22922 | 20.84 | 1.1412
LS1 | Beta .085 QWR | #10 | TX-881B | Before Solenoid 1 | x111782 | A5 | 83.6988 | 97.5Ω, OK | 104.85 | 83.94 |  | 0.24122 | 20.91 | 0.67078
LS1 | Beta .085 QWR | #10 | TX-882A | After Solenoid 1 | x111643 | A5 | 85.5847 | 87.9Ω, OK | 106.24 | 85.78 |  | 0.19527 | 20.46 | 1.891
LS1 | Beta .085 QWR | #10 | TX-882B | After Solenoid 1 | x111668 | A5 | 83.889 | 84.7Ω, OK | 104.57 | 84.1 |  | 0.21101 | 20.47 | 1.4847
LS1 | Beta .085 QWR | #10 | TX-883A | Before Solenoid 2 | x111722 | A5 | 85.7062 | 86.4Ω, OK | 106.64 | 85.93 |  | 0.22375 | 20.71 | 0.91857
LS1 | Beta .085 QWR | #10 | TX-883B | Before Solenoid 2 | x111716 | A5 | 85.0114 | 85.5Ω, OK | 105.49 | 85.23 |  | 0.21857 | 20.26 | 0.47625
LS1 | Beta .085 QWR | #10 | TX-884A | After Solenoid 2 | x111758 | A5 | 85.1409 | 85.4Ω, OK | 104.91 | 85.36 |  | 0.21915 | 19.55 | 1.3699
LS1 | Beta .085 QWR | #10 | TX-884B | After Solenoid 2 | x111789 | A5 | 83.9901 | 84.2Ω, OK | 104.34 | 84.21 |  | 0.21987 | 20.13 | 0.93085
LS1 | Beta .085 QWR | #10 | TX-885A | Before Solenoid 3 | x111771 | A6 | 85.8647 | 86.7Ω, OK | 106.95 | 86.08 |  | 0.21533 | 20.87 | 2.0879
LS1 | Beta .085 QWR | #10 | TX-885B | Before Solenoid3 | x111778 | A6 | 83.9921 | 97.4Ω, OK | 104.88 | 84.18 |  | 0.18791 | 20.7 | 1.6847
LS1 | Beta .085 QWR | #10 | TX-886A | After Solenoid 3 | x111644 | A6 | 84.075 | 85.0Ω, OK | 106.11 | 84.27 |  | 0.19497 | 21.84 | 0.46659
LS1 | Beta .085 QWR | #10 | TX-886B | After Solenoid 3 | x111721 | A6 | 83.8034 | 84.8Ω, OK | 105.13 | 83.99 |  | 0.18659 | 21.14 | 0.08503
LS1 | Beta .085 QWR | #11 | TX-881A | Before Solenoid 1 | x111742 | A5 | 85.231 | 85.1 Ω, OK | 106.408 | 85.441 |  | 0.21005 | 20.967 | 0.16467
LS1 | Beta .085 QWR | #11 | TX-881B | Before Solenoid 1 | x111773 | A5 | 85.2763 | 94.0 Ω, OK | 106.131 | 85.506 |  | 0.22967 | 20.625 | 0.27505
LS1 | Beta .085 QWR | #11 | TX-882A | After Solenoid 1 | x111720 | A5 | 84.382 | 84.2 Ω, OK | 105.065 | 84.6 |  | 0.21799 | 20.465 | 1.0811
LS1 | Beta .085 QWR | #11 | TX-882B | After Solenoid 1 | x111781 | A5 | 85.6811 | 85.7 Ω, OK | 106.276 | 85.897 |  | 0.21593 | 20.379 | 1.515
LS1 | Beta .085 QWR | #11 | TX-883A | Before Solenoid 2 | x111714 | A5 | 85.6069 | 85.4 Ω, OK | 105.984 | 85.819 |  | 0.2121 | 20.165 | 2.0605
LS1 | Beta .085 QWR | #11 | TX-883B | Before Solenoid 2 | x111679 | A5 | 83.7585 | 83.5 Ω, OK | 104.132 | 83.957 |  | 0.19847 | 20.175 | 1.6499
LS1 | Beta .085 QWR | #11 | TX-884A | After Solenoid 2 | x111684 | A5 | 84.5424 | 84.4 Ω, OK | 104.609 | 84.76 |  | 0.21758 | 19.849 | 1.2247
LS1 | Beta .085 QWR | #11 | TX-884B | After Solenoid 2 | x111689 | A5 | 85.9847 | 85.8 Ω, OK | 106.57 | 86.198 |  | 0.21328 | 20.372 | 1.6556
LS1 | Beta .085 QWR | #11 | TX-885A | Before Solenoid 3 | x111678 | A6 | 85.8764 | 85.7 Ω, OK | 106.582 | 86.079 |  | 0.20265 | 20.503 | 2.0754
LS1 | Beta .085 QWR | #11 | TX-885B | Before Solenoid3 | x111671 | A6 | 84.0036 | 92.5 Ω, OK | 104.814 | 84.176 |  | 0.17239 | 20.638 | 1.7004
LS1 | Beta .085 QWR | #11 | TX-886A | After Solenoid 3 | x111646 | A6 | 84.9751 | 84.9 Ω, OK | 106.347 | 85.183 |  | 0.20792 | 21.164 | 1.7704
LS1 | Beta .085 QWR | #11 | TX-886B | After Solenoid 3 | x111674 | A6 | 83.4126 | 83.5 Ω, OK | 104.601 | 83.614 |  | 0.20141 | 20.987 | 1.3611
LS2 | Beta .29 HWR | #2 | TX-281A | Before Solenoid 1 | x111759 | A5 | 84.5366 | 85.2 Ω,ok | 104.427 | 84.503 |  | 0.03356 | 19.924
LS2 | Beta .29 HWR | #2 | TX-281B | Before Solenoid 1 | x111768 | A5 | 86.4223 | 89.0 Ω,ok | 106.493 | 86.371 |  | 0.05129 | 20.122
LS2 | Beta .29 HWR | #2 | TX-282A | After Solenoid 1 | x111677 | A5 | 86.092 | 86.6 Ω,ok | 105.986 | 86.049 |  | 0.04303 | 19.937
LS2 | Beta .29 HWR | #2 | TX-282B | After Solenoid 1 | x113495 | A5 | 55.3134 | 55.7 Ω,ok | 75.133 | 55.275 |  | 0.03841 | 19.858
LS2 | Beta .29 HWR | #3 | TX-281A | Before Solenoid 1 | x111774 | A5 | 84.385 | 84.8 | 104.364 | 84.33 |  | 0.05503 | 20.034
LS2 | Beta .29 HWR | #3 | TX-281B | Before Solenoid 1 | x111788 | A5 | 84.61 | 85.1 | 104.07 | 84.494 |  | 0.116 | 19.576
LS2 | Beta .29 HWR | #3 | TX-282A | After Solenoid 1 | x111675 | A5 | 86.1088 | 86.1 | 105.855 | 86.044 |  | 0.0648 | 19.811
LS2 | Beta .29 HWR | #3 | TX-282B | After Solenoid 1 | x111745 | A5 | 84.2801 | 84.8 | 104.32 | 84.201 |  | 0.07905 | 20.119
LS2 | Beta .29 HWR | #4 | TX-281A | Before Solenoid 1 | x111785 | A5 | 84.26 | 84.8 | 104.77 | 84.588
LS2 | Beta .29 HWR | #4 | TX-281B | Before Solenoid 1 | x111727 | A5 | 83.88 | 84.4 | 104.498 | 84.187
LS2 | Beta .29 HWR | #4 | TX-282A | After Solenoid 1 | x111725 | A5 | 84.39 | 84.5 | 105.21 | 84.697
LS2 | Beta .29 HWR | #4 | TX-282B | After Solenoid 1 | x111728 | A5 | 86.03 | 86.4 | 106.582 | 86.354
LS2 | Beta .29 HWR | #5 | TX-281A | Before Solenoid 1 | x111682 | A5 | 83.8122 | 83.8 | 103.583 | 84.037
LS2 | Beta .29 HWR | #5 | TX-281B | Before Solenoid 1 | x111641 | A5 | 85.1023 | 85.1 | 105.193 | 85.303
LS2 | Beta .29 HWR | #5 | TX-282A | After Solenoid 1 | x111685 | A5 | 86.0957 | 86.2 | 106.47 | 86.335
LS2 | Beta .29 HWR | #5 | TX-282B | After Solenoid 1 | x111783 | A5 | 84.771 | 84.9 | 105.239 | 85.016
LS2 | Beta .29 HWR | #6 | TX-281A | Before Solenoid 1 | x111715 | A5 | 83.8623 | 84.2 | 104.06 | 83.892
LS2 | Beta .29 HWR | #6 | TX-281B | Before Solenoid 1 | x111681 | A5 | 84.4967 | 84.7 | 104.554 | 84.365
LS2 | Beta .29 HWR | #6 | TX-282A | After Solenoid 1 | x111756 | A5 | 86.8786 | 86.9 | 106.764 | 86.634
LS2 | Beta .29 HWR | #6 | TX-282B | After Solenoid 1 | x110680 | A5 | 89.9265 | 90.1 | 109.571 | 89.841
LS2 | Beta .29 HWR | #7 | TX-281A | Before Solenoid 1 | x111690 | A5 | 85.8432 | 85.5
LS2 | Beta .29 HWR | #7 | TX-281B | Before Solenoid 1 | x111691 | A5 | 85.9542 | 85.7
LS2 | Beta .29 HWR | #7 | TX-282A | After Solenoid 1 | x113742 | A5 | 55.1145 | 55.4
LS2 | Beta .29 HWR | #7 | TX-282B | After Solenoid 1 | x113760 | A5 | 56.6671 | 56.6
LS2 | Beta .29 HWR | #8 | TX-281A | Before Solenoid 1 | x113520 | A5 | 55.3248 | 55 | 75.13 | 55.092
LS2 | Beta .29 HWR | #8 | TX-281B | Before Solenoid 1 | x113523 | A5 | 54.5217 | 54.2 | 74.234 | 54.299
LS2 | Beta .29 HWR | #8 | TX-282A | After Solenoid 1 | x113528 | A5 | 52.8702 | 52.5 | 72.627 | 52.654
LS2 | Beta .29 HWR | #8 | TX-282B | After Solenoid 1 | x113522 | A5 | 57.0874 | 56.7 | 76.702 | 56.859
LS2 | Beta .29 HWR | #9 | TX-281A | Before Solenoid 1 | x113524 | A5 | 52.821 | 52.7 | 72.475 | 52.678
LS2 | Beta .29 HWR | #9 | TX-281B | Before Solenoid 1 | x113521 | A5 | 54.0097 | 53.9 | 73.623 | 53.874
LS2 | Beta .29 HWR | #9 | TX-282A | After Solenoid 1 | x113518 | A5 | 52.8104 | 52.8 | 72.55 | 52.672
LS2 | Beta .29 HWR | #9 | TX-282B | After Solenoid 1 | x113516 | A5 | 62.45 | 62.5 | 82.3 | 62.35
LS2 | Beta .29 HWR | #10 | TX-281A | Before Solenoid 1 | x111757 | A5 | 84.2006 | 83.9 | 104.23 | 83.921
LS2 | Beta .29 HWR | #10 | TX-281B | Before Solenoid 1 | x111790 | A5 | 83.8764 | 83.6 | 103.52 | 83.6
LS2 | Beta .29 HWR | #10 | TX-282A | After Solenoid 1 | x111686 | A5 | 85.2919 | 85 | 105.08 | 84.993
LS2 | Beta .29 HWR | #10 | TX-282B | After Solenoid 1 | x111776 | A5 | 85.2344 | 84.9 | 105.39 | 84.933
LS2 | Beta .29 HWR | #11 | TX-281A | Before Solenoid 1 | x111748 | A5 | 86.22 | 84.9 | 105.538 | 85.438
LS2 | Beta .29 HWR | #11 | TX-281B | Before Solenoid 1 | x111760 | A5 | 83.55 | 82.3 | 102.822 | 82.82
LS2 | Beta .29 HWR | #11 | TX-282A | After Solenoid 1 | x111670 | A5 | 84.1 | 82.9 | 103.653 | 83.385
LS2 | Beta .29 HWR | #11 | TX-282B | After Solenoid 1 | x111786 | A5 | 83.87 | 82.7 | 102.486 | 83.186
LS2 | Beta .29 HWR | #12 | TX-281A | Before Solenoid 1 | x111743 | A5 | 86.2905 | 86.2
LS2 | Beta .29 HWR | #12 | TX-281B | Before Solenoid 1 | x111645 | A5 | 84.58 | 84.4
LS2 | Beta .29 HWR | #12 | TX-282A | After Solenoid 1 | x111749 | A5 | 82.8966 | 82.9
LS2 | Beta .29 HWR | #12 | TX-282B | After Solenoid 1 | x111767 | A5 | 85.0437 | 84.9
LS1 | Beta .085 QWR | #12 | TX-881A | Before Solenoid 1 | x113738 | A5 | 54.78 | 54.9 Ω, OK | 75.565 | 54.677
LS1 | Beta .085 QWR | #12 | TX-881B | Before Solenoid 1 | x113710 | A5 | 57.6569 | 57.7 Ω, OK | 78.66 | 57.521
LS1 | Beta .085 QWR | #12 | TX-882A | After Solenoid 1 | x113816 | A5 | 59.1895 | 59.1 Ω, OK | 79.158 | 58.965
LS1 | Beta .085 QWR | #12 | TX-882B | After Solenoid 1 | x113754 | A5 | 59.5311 | 59.5 Ω, OK | 79.58 | 59.363
LS1 | Beta .085 QWR | #12 | TX-883A | Before Solenoid 2 | x113746 | A5 | 55.1295 | 55.1 Ω, OK | 74.957 | 55.012
LS1 | Beta .085 QWR | #12 | TX-883B | Before Solenoid 2 | x113504 | A5 | 62.1356 | 62.1 Ω, OK | 82.028 | 62.03
LS1 | Beta .085 QWR | #12 | TX-884A | After Solenoid 2 | x113820 | A5 | 58.9231 | 58.7 Ω, OK | 78.786 | 58.759
LS1 | Beta .085 QWR | #12 | TX-884B | After Solenoid 2 | x113503 | A5 | 61.7911 | 61.5 Ω, OK | 81.594 | 61.567
LS1 | Beta .085 QWR | #12 | TX-885A | Before Solenoid 3 | x113751 | A6 | 57.5281 | 57.5 Ω, OK | 77.487 | 57.34
LS1 | Beta .085 QWR | #12 | TX-885B | Before Solenoid 3 | x113747 | A6 | 57.3114 | 57.3 Ω, OK | 77.256 | 57.184
LS1 | Beta .085 QWR | #12 | TX-886A | After Solenoid 3 | x113723 | A6 | 55.4193 | 55.3 Ω, OK | 76.281 | 55.21
LS1 | Beta .085 QWR | #12 | TX-886B | After Solenoid 3 | x113800 | A6 | 56.2635 | 56.2 Ω, OK | 77.001 | 56.083

## Sensor Testing Data
Cernox 1050 Measurements | [7 empty] | Installation Data
Batch | Test Number | Serial Number | Resistance at 290K=62.33F | 295K=71.33F | 300K=80.33F | Tested Resistance | Notes | installed in cyro module# | installation location | device tag | feedthrough port | installed resistance testing
1 | 1 | X110635 | 86.9715 | 85.5574 | 84.1897 | 85.2908 |  | B041 Cyromodule #1 | After Solenoid 3 | TX-485A | A6 | 106.7,ok
| 2 | X110648 | 94.3666 | 92.8743 | 91.4305 | 92.4949 |  | B041 Cyromodule #1 | After Solenoid 3 | TX-486A | A6 | 104.5,ok
| 3 | X110649 | 87.6981 | 86.2761 | 84.9007 | 86.4785
| 4 | X110650 | 86.9905 | 85.5664 | 84.1891 | 85.7023 |  | B085 Cyromodule #3 | Before Solenoid 1 | TX-881A | A5 | 94.1,ok
| 5 | X110651 | 86.9159 | 85.5139 | 84.1576 | 85.2382 |  | B085 Cyromodule #3 | Before Solenoid 2 | TX-883A | A5 | 94.2,ok
| 6 | X110652 | 94.4894 | 92.9825 | 91.5248 | 93.2228 |  | B085 Cyromodule #3 | Before Solenoid3 | TX-885B | A6 | 101.6,ok
| 7 | X110653 | 86.6632 | 85.2945 | 83.97 | 84.9265 |  | B085 Cyromodule #3 | After Solenoid 2 | TX-884B | A5 | 93.9,ok
| 8 | X110664 | 85.144 | 83.7949 | 82.4895 | 83.7174 |  | B085 Cyromodule #3 | After Solenoid 2 | TX-884A | A5 | 92.6,ok
| 9 | X110665 | 89.2433 | 87.8243 | 86.4515 | 87.9623 |  | B085 Cyromodule #3 | Before Solenoid 3 | TX-885A | A6 | 96.5,ok
| 10 | X110666 | 87.4389 | 86.0289 | 84.6649 | 85.8725 |  | B085 Cyromodule #3 | After Solenoid 3 | TX-886B | A6 | 94.8,ok
| 11 | X110667 | 86.6012 | 85.1814 | 83.8083 | 84.9701 |  | B041 Cyromodule #1 | Before Solenoid 1 | TX-481B | A5 | 105.1,ok
| 12 | X110668 | 89.8348 | 88.3951 | 87.0024 | 88.244 |  | B041 Cyromodule #1 | After Solenoid 2 | TX-484A | A5 | 105.9,ok
| 13 | X110669 | 87.4688 | 86.0642 | 84.7054 | 85.5637 |  | B041 Cyromodule #1 | Before Solenoid 2 | TX-483A | A5 | 102.4,ok
| 14 | X110670 | 88.6944 | 87.2669 | 85.8862 | 87.0374 |  | B041 Cyromodule #1 | After Solenoid 2 | TX-484B | A5 | 106.9,ok
| 15 | X110671 | 85.4981 | 84.1293 | 82.8049 | 83.8912 |  | B041 Cyromodule #1 | Before Solenoid 2 | TX-483B | A5 | 104.5,ok
| 16 | X110672 | 87.0253 | 85.6207 | 84.262 | 85.3671 |  | B041 Cyromodule #1 | After Solenoid 1 | TX-482A | A5 | 114.2,ok
| 17 | X110673 | 97.598 | 96.0351 | 94.5233 | 95.732 |  | B041 Cyromodule #1 | After Solenoid 1 | TX-482B | A5 | 103.7,ok
| 18 | X110676 | 88.3863 | 86.9811 | 85.6216 | 86.5405 |  | B041 Cyromodule #1 | Before Solenoid 1 | TX-481A | A5 | 103.4,ok
| 19 | X110677 | 87.0499 | 85.6261 | 84.2492 | 85.764 |  | B041 Cyromodule #1 | After Solenoid 3 | TX-486B | A6 | 111.6,ok
| 20 | X110678 | 89.515 | 88.0997 | 86.7304 | 87.8621 |  | B041 Cyromodule #1 | After Solenoid 3 | TX-485B | A6 | 104.3,ok
2 | 21 | X110680 | 91.5945 | 90.1262 | 88.7059 | 89.9265 |  | B29 Cyromodule #6 | After Solenoid 1 | TX-282B | A5 | 90.1
| 22 | X111641 | 86.5761 | 85.1023 | 83.6818 | 84.5376 |  | B29 Cyromodule #5 | Before Solenoid 1 | TX-281B |  | 85.1
| 23 | X111643 | 87.0721 | 85.5847 | 84.1512 | 85.5028 |  | B085 Cyromodule #10 | After Solenoid 1 | TX-882A | A5 | 87.9Ω, OK
| 24 | X111644 | 85.5167 | 84.075 | 82.6852 | 84.0195 |  | B085 Cyromodule #10 | After Solenoid 3 | TX-886A | A6 | 85.0Ω, OK
| 25 | X111645 | 86.0314 | 84.58 | 83.1808 | 84.5949 |  | B29 Cyromodule #12 | Before Solenoid 1 | TX-281B | A5 | 84.4
| 26 | X111646 | 86.4389 | 84.9751 | 83.564 | 84.9592 |  | B085 Cyromodule #11 | After Solenoid 3 | TX-886A | A6 | 84.9 Ω, OK
| 27 | X111668 | 85.3337 | 83.889 | 82.4963 | 84.0994 |  | B085 Cyromodule #10 | After Solenoid 1 | TX-882B | A5 | 84.7Ω, OK
| 28 | X111670 | 85.5422 | 84.0969 | 82.7036 | 84.2168 |  | B29 Cyromodule #11 | After Solenoid 1 | TX-282A | A5 | 82.9
| 29 | X111671 | 85.4427 | 84.0036 | 82.6162 | 83.9857 |  | B085 Cyromodule #11 | Before Solenoid3 | TX-885B | A6 | 92.5 Ω, OK
| 30 | X111674 | 84.8354 | 83.4126 | 82.0407 | 83.4743 |  | B085 Cyromodule #11 | After Solenoid 3 | TX-886B | A6 | 83.5 Ω, OK
| 31 | X111675 | 87.6038 | 86.1088 | 84.6679 | 85.631
| 32 | X111677 | 87.5927 | 86.092 | 84.6457 | 86.2328 |  | B29 Cyromodule #2 | After Solenoid 1 | TX-282A |  | 86.6 Ω,ok
| 33 | X111678 | 87.3662 | 85.8764 | 84.4404 | 85.6554 |  | B085 Cyromodule #11 | Before Solenoid 3 | TX-885A | A6 | 85.7 Ω, OK
| 34 | X111679 | 85.1993 | 83.7585 | 82.3696 | 83.8087 |  | B085 Cyromodule #11 | Before Solenoid 2 | TX-883B | A5 | 83.5 Ω, OK
| 35 | X111681 | 85.9678 | 84.51 | 83.1047 | 84.4967 |  | B29 Cyromodule #6 | Before Solenoid 1 | TX-281B | A5 | 84.7
| 36 | X111682 | 85.2576 | 83.8122 | 82.419 | 83.9905 |  | B29 Cyromodule #5 | Before Solenoid 1 | TX-281A |  | 83.8
| 37 | X111684 | 86.0029 | 84.5424 | 83.1346 | 84.6192 |  | B085 Cyromodule #11 | After Solenoid 2 | TX-884A | A5 | 84.4 Ω, OK
| 38 | X111685 | 87.5967 | 86.0957 | 84.6491 | 86.1692 |  | B29 Cyromodule #5 | After Solenoid 1 | TX-282A |  | 86.1
| 39 | X111686 | 86.7706 | 85.2919 | 83.8667 | 85.1765 |  | B29 Cyromodule #10 | After Solenoid 1 | TX-282A |  | 85
| 40 | X111689 | 87.4804 | 85.9847 | 84.5433 | 85.8682 |  | B085 Cyromodule #11 | After Solenoid 2 | TX-884B | A5 | 85.8 Ω, OK
| 41 | X111690 | 87.2486 | 85.7578 | 84.3209 | 85.8432
| 42 | X111691 | 87.4038 | 85.9068 | 84.464 | 85.9542
| 43 | X111714 | 87.094 | 85.6069 | 84.1736 | 85.4662 |  | B085 Cyromodule #11 | Before Solenoid 2 | TX-883A | A5 | 85.4 Ω, OK
| 44 | X111715 | 85.4865 | 84.0446 | 82.6545 | 83.8623 |  | B29 Cyromodule #6 | Before Solenoid 1 | TX-281A | A5 | 84.2
| 45 | X111716 | 86.4842 | 85.0114 | 83.5919 | 85.0245 |  | B085 Cyromodule #10 | Before Solenoid 2 | TX-883B | A5 | 85.5Ω, OK
| 46 | X111719 | 87.3101 | 85.8165 | 84.377 | 85.1815
| 47 | X111720 | 85.8395 | 84.382 | 82.9771 | 84.2932 |  | B085 Cyromodule #11 | After Solenoid 1 | TX-882A | A5 | 84.2 Ω, OK
| 48 | X111721 | 85.2221 | 83.8034 | 82.4354 | 83.0019 |  | B085 Cyromodule #10 | After Solenoid 3 | TX-886B | A6 | 84.8Ω, OK
| 49 | X111722 | 87.1941 | 85.7062 | 84.2723 | 85.7057 |  | B085 Cyromodule #10 | Before Solenoid 2 | TX-883A | A5 | 86.4Ω, OK
| 50 | X111725 | 85.8448 | 84.3893 | 82.9863 | 84.3575
| 51 | X111727 | 85.3186 | 83.8795 | 82.4921 | 84.4089
| 52 | X111728 | 87.5248 | 86.0316 | 84.5925 | 86.0294
| 53 | X111742 | 86.7037 | 85.231 | 83.8113 | 85.1972 |  | B085 Cyromodule #11 | Before Solenoid 1 | TX-881A | A5 | 85.1 Ω, OK
| 54 | X111743 | 87.7875 | 86.2905 | 84.8478 | 85.7505 |  | B29 Cyromodule #12 | Before Solenoid 1 | TX-281A | A5 | 86.2
| 55 | X111745 | 85.7291 | 84.2801 | 82.8832 | 83.8894
| 56 | X111748 | 87.7263 | 86.2231 | 84.7744 | 86.2305 |  | B29 Cyromodule #11 | Before Solenoid 1 | TX-281A | A5 | 84.9
| 57 | X111749 | 84.3029 | 82.8966 | 81.5407 | 82.3185 |  | B29 Cyromodule #12 | After Solenoid 1 | TX-282A | A5 | 82.9
| 58 | X111756 | 88.3862 | 86.8718 | 85.4123 | 86.8786 |  | B29 Cyromodule #6 | After Solenoid 1 | TX-282A | A5 | 86.9
| 59 | X111757 | 85.6508 | 84.2006 | 82.8026 | 84.4622 |  | B29 Cyromodule #10 | Before Solenoid 1 | TX-281A | A5 | 83.9
| 60 | X111758 | 86.6129 | 85.1409 | 83.7219 | 85.1622 |  | B085 Cyromodule #10 | After Solenoid 2 | TX-884A | A5 | 85.4Ω, OK
| 61 | X111759 | 85.9972 | 84.5366 | 83.1287 | 84.5259 |  | B29 Cyromodule #2 | Before Solenoid 1 | TX-281A |  | 85.2 Ω,ok
| 62 | X111760 | 84.9731 | 83.5491 | 82.1762 | 83.3235 |  | B29 Cyromodule #11 | Before Solenoid 1 | TX-281B | A5 | 82.3
| 63 | X111767 | 86.5146 | 85.0437 | 83.6259 | 84.8849 |  | B29 Cyromodule #12 | After Solenoid 1 | TX-282B | A5 | 84.9
| 64 | X111768 | 87.9232 | 86.4223 | 84.9758 | 86.4405 |  | B29 Cyromodule #2 | Before Solenoid 1 | TX-281B |  | 89.0 Ω,ok
| 65 | X111769 | 86.0665 | 84.6108 | 83.2075 | 84.7972 |  | B085 Cyromodule #10 | Before Solenoid 1 | TX-881A | A5 | 99.6Ω, OK
| 66 | X111771 | 87.3552 | 85.8647 | 84.4281 | 86.0405 |  | B085 Cyromodule #10 | Before Solenoid 3 | TX-885A | A6 | 86.7Ω, OK
| 67 | X111773 | 86.7542 | 85.2763 | 83.8519 | 85.0984 |  | B085 Cyromodule #11 | Before Solenoid 1 | TX-881B | A5 | 94.0 Ω, OK
| 68 | X111774 | 85.8385 | 84.385 | 82.9839 | 84.6804
| 69 | X111776 | 86.7119 | 85.2344 | 83.8103 | 85.4386 |  | B29 Cyromodule #10 | After Solenoid 1 | TX-282B |  | 84.9
| 70 | X111777 | 86.5208 | 85.0523 | 83.6368 | 84.9297
| 71 | X111778 | 85.4185 | 83.9921 | 82.6167 | 84.0444 |  | B085 Cyromodule #10 | Before Solenoid3 | TX-885B | A6 | 97.4Ω, OK
| 72 | X111781 | 87.1651 | 85.6811 | 84.2507 | 85.7536 |  | B085 Cyromodule #11 | After Solenoid 1 | TX-882B | A5 | 85.7 Ω, OK
| 73 | X111782 | 85.1361 | 83.6988 | 82.3131 | 83.7476 |  | B085 Cyromodule #10 | Before Solenoid 1 | TX-881B | A5 | 97.5Ω, OK
| 74 | X111783 | 86.2378 | 84.771 | 83.3571 | 84.5378 |  | B29 Cyromodule #5 | After Solenoid 1 | TX-282B |  | 84.9
| 75 | X111784 | 85.3493 | 83.911 | 82.5244 | 83.5747
| 76 | X111785 | 85.7139 | 84.262 | 82.8624 | 84.3403
| 77 | X111786 | 85.3078 | 83.8681 | 82.4801 | 83.9487 |  | B29 Cyromodule #11 | After Solenoid 1 | TX-282B | A5 | 82.7
| 78 | X111788 | 86.0744 | 84.6119 | 83.2021 | 83.9561
| 79 | X111789 | 85.4337 | 83.9901 | 82.5984 | 83.9875 |  | B085 Cyromodule #10 | After Solenoid 2 | TX-884B | A5 | 84.2Ω, OK
| 80 | X111790 | 85.3159 | 83.8764 | 82.4886 | 83.8972 |  | B29 Cyromodule #10 | Before Solenoid 1 | TX-281B | A5 | 83.6
3 | 81 | X110679 | 87.2389 | 85.8282 | 84.4639 | 85.9123 |  | B085 Cyromodule#6 | Before Solenoid 2 | TX-883B | A5 | 105.6 Ω, OK
| 82 | X112276 | 74.8453 | 73.5645 | 72.3298 | 73.4285 |  | B085 Cyromodule#6 | After Solenoid 1 | TX-882B | A5 | 94.0 Ω, OK
| 83 | X112310 | 74.9881 | 73.7009 | 72.4602 | 73.8146
| 84 | X112343 | 79.0955 | 77.7049 | 76.3651 | 77.8495 |  | B085 Cyromodule#6 | After Solenoid 2 | TX-884A | A5 | 98.0 Ω, OK
| 85 | X112344 | 73.8684 | 72.6156 | 71.4074 | 72.7809
| 86 | X112345 | 80.2373 | 78.8203 | 77.4553 | 78.8445 |  | B085 Cyromodule#6 | Before Solenoid3 | TX-885B | A6 | 99.4 Ω, OK
| 87 | X112346 | 71.9125 | 70.711 | 69.552 | 70.7015 |  | B085 Cyromodule#6 | After Solenoid 1 | TX-882A | A5 | 91.1 Ω, OK
| 88 | X112347 | 75.7977 | 74.4938 | 73.2368 | 74.2795 |  | B085 Cyromodule#6 | Before Solenoid 1 | TX-881B | A5 | 95.1 Ω, OK
| 89 | X112348 | 78.0742 | 76.7115 | 75.3984 | 76.4448 |  | B085 Cyromodule#6 | Before Solenoid 1 | TX-881A | A5 | 97.4 Ω, OK
| 90 | X112349 | 75.8641 | 74.5603 | 73.3035 | 74.4392 |  | B085 Cyromodule#6 | Before Solenoid 3 | TX-885A | A6 | 95.0 Ω, OK
| 91 | X112350 | 70.1066 | 68.9416 | 67.8179 | 68.9115 |  | B085 Cyromodule#6 | Before Solenoid 2 | TX-883A | A5 | 88.6 Ω, OK
| 92 | X112351 | 79.767 | 78.363 | 77.0104 | 78.2997 |  | B085 Cyromodule#6 | After Solenoid 2 | TX-884B | A5 | 98.2 Ω, OK
| 93 | X112352 | 78.4486 | 77.0748 | 75.7512 | 67.2152 | Out of range - Replaced by Lakeshore (see R1 batch)
| 94 | X112353 | 72.8689 | 71.7364 | 70.6439 | 76.4151 | Out of range - Replaced by Lakeshore (see R1 batch)
| 95 | X112355 | 76.2311 | 74.9136 | 73.6438 | 74.5348 | No Polarity
| 96 | X112356 | 79.5791 | 78.1917 | 76.8547 | 78.1925 |  | B085 Cyromodule #8 | After Solenoid 3 | TX-886A | A6 | 98.3, ok
| 97 | X112359 | 73.0809 | 71.8461 | 70.6554 | 71.4461
| 98 | X112361 | 82.2488 | 80.7918 | 79.3881 | 81.0859
| 99 | X112362 | 80.7058 | 79.2797 | 77.9059 | 79.3903
| 100 | X112363 | 77.8586 | 76.499 | 75.1889 | 75.9848
| 101 | X112368 | 78.2073 | 76.8417 | 75.526 | 76.8065
| 102 | X112369 | 73.224 | 71.9907 | 70.8013 | 72.1192
| 103 | X112370 | 71.0849 | 69.9033 | 68.7635 | 69.9506
| 104 | X112371 | 74.871 | 73.585 | 72.3455 | 73.5665
| 105 | X112372 | 79 | 77.6146 | 76.2797 | 77.5525
| 106 | X112373 | 65.6251 | 64.5538 | 63.5203 | 64.5486 |  | B041 Cyromodule#2 | After Solenoid 1 | TX-482B | A5 | 83.6,ok
| 107 | X112375 | 77.9281 | 76.5678 | 75.2571 | 75.8205 |  | B041 Cyromodule#3 | After Solenoid 1 | TX-482A | A5 | 95.7,ok
| 108 | X112377 | 78.0086 | 76.6471 | 75.335 | 76.2645 |  | B041 Cyromodule#2 | After Solenoid 2 | TX-484A | A5 | 95.2,ok
| 109 | X112379 | 67.5261 | 66.4566 | 65.4246 | 66.3005 |  | B041 Cyromodule#2 | Before Solenoid 2 | TX-483B | A5 | 84.9,ok
| 110 | X112388 | 75.1642 | 73.8723 | 72.6271 | 73.2664 |  | B041 Cyromodule#2 | After Solenoid 3 | TX-486A | A6 | 93.4,ok
| 111 | X112389 | 77.0098 | 75.6705 | 74.3798 | 74.7554 |  | B041 Cyromodule#2 | After Solenoid 3 | TX-485A | A6 | 94.7,ok
| 112 | X112390 | 75.7811 | 74.4768 | 73.2195 | 74.0931 |  | B041 Cyromodule#3 | After Solenoid 1 | TX-482B | A5 | 93.2,ok
| 113 | X112391 | 83.4963 | 82.0006 | 80.5598 | 81.1253 |  | B041 Cyromodule#3 | Before Solenoid 1 | TX-481B | A5 | 100.8,ok
| 114 | X112392 | 76.5483 | 75.2203 | 73.9406 | 75.1612 |  | B041 Cyromodule#2 | After Solenoid 3 | TX-485B | A6 | 94.4,ok
| 115 | X112393 | 76.4185 | 75.0939 | 73.8173 | 75.1578 |  | B041 Cyromodule#3 | After Solenoid 2 | TX-484B | A5 | 94.5,ok
| 116 | X112401 | 67.9573 | 66.841 | 65.7641 | 66.7642 |  | B085 Cyromodule #3 | After Solenoid 1 | TX-882A | A5 | 75.5,ok
| 117 | X112403 | 82.4494 | 80.9873 | 79.5787 | 80.0756 |  | B085 Cyromodule #8 | After Solenoid 3 | TX-886B | A6 | 101.3, ok
| 118 | X112406 | 73.7116 | 72.4607 | 71.2546 | 72.5852 |  | B041 Cyromodule#3 | Before Solenoid 2 | TX-483B | A5 | 91.2,ok
| 119 | X112408 | 76.4939 | 75.1693 | 73.8928 | 75.2302 |  | B041 Cyromodule#3 | Before Solenoid 1 | TX-481A | A5 | 94.5,ok
| 120 | X112409 | 79.5595 | 78.1543 | 76.8005 | 78.2356 |  | B041 Cyromodule#3 | Before Solenoid 2 | TX-483A | A5 | 96.9,ok
| 121 | X112412 | 76.1035 | 74.7875 | 73.5192 | 74.5119 |  | B041 Cyromodule#3 | After Solenoid 3 | TX-485A | A6 | 94.5,ok
| 122 | X112413 | 77.2494 | 75.9064 | 74.6121 | 75.8189 |  | B041 Cyromodule#3 | After Solenoid 3 | TX-486A | A6 | 95.1,ok
| 123 | X112415 | 67.4371 | 66.3334 | 65.2685 | 66.3954 | No Polarity | B085 Cyromodule #8 | Before Solenoid 3 | TX-885B | A6 | 85.9, ok
| 124 | X112416 | 77.5964 | 76.2489 | 74.95 | 76.1618 |  | B085 Cyromodule #3 | After Solenoid 3 | TX-886A | A6 | 94.9,ok
| 125 | X112417 | 82.4819 | 81.0172 | 79.606 | 80.8677 |  | B041 Cyromodule#3 | After Solenoid 2 | TX-484A | A5 | 100.3,ok
| 126 | X112419 | 78.8135 | 77.4327 | 76.1022 | 77.2074 |  | B041 Cyromodule#2 | Before Solenoid 1 | TX-481A | A5 | 95.8,ok
| 127 | X112420 | 71.655 | 70.4426 | 69.2737 | 70.1982
| 128 | X112421 | 80.444 | 79.02 | 77.6482 | 78.9899 |  | B085 Cyromodule#6 | After Solenoid 3 | TX-886B | A6 | 99.0 Ω, OK
| 129 | X112422 | 76.061 | 74.7496 | 73.4857 | 74.6323 |  | B085 Cyromodule#6 | After Solenoid 3 | TX-886A | A6 | 95.3 Ω, OK
| 130 | X112424 | 71.7716 | 70.5614 | 69.3946 | 70.4238 |  | B085 Cyromodule#7 | After Solenoid 1 | TX-882A | A5 | 90.5 Ω, OK
| 131 | X112425 | 76.291 | 74.9701 | 73.6971 | 74.8864 |  | B085 Cyromodule#7 | Before Solenoid 1 | TX-881A | A5 | 96.0 Ω, OK
| 132 | X112426 | 75.4588 | 74.1722 | 72.9317 | 74.0147 |  | B085 Cyromodule#7 | Before Solenoid 1 | TX-881B | A5 | 95.3 Ω, OK
| 133 | X112427 | 80.7498 | 79.3307 | 77.9632 | 78.3454 |  | B041 Cyromodule#2 | After Solenoid 3 | TX-486B | A6 | 98.3,ok
| 134 | X112428 | 81.4544 | 80.0162 | 78.6305 | 79.9781 |  | B041 Cyromodule#3 | After Solenoid 3 | TX-485B | A6 | 99.5,ok
| 135 | X112429 | 77.17 | 75.8334 | 74.5452 | 75.6947 |  | B041 Cyromodule#3 | After Solenoid 3 | TX-486B | A6 | 94.8,ok
| 136 | X112430 | 81.215 | 79.7837 | 78.4045 | 78.9161 |  | B041 Cyromodule#2 | Before Solenoid 2 | TX-483A | A5 | 98.1,ok
| 137 | X112431 | 67.7871 | 66.7422 | 65.7334 | 66.6084 |  | B041 Cyromodule#2 | After Solenoid 2 | TX-484B | A5 | 85.5,ok
| 138 | X112432 | 68.0264 | 66.9773 | 65.9643 | 66.4855 |  | B041 Cyromodule#2 | Before Solenoid 1 | TX-481B | A5 | 85.4,ok
| 139 | X112433 | 81.1557 | 79.7267 | 78.3502 | 79.3089
| 140 | X112434 | 67.3944 | 66.3595 | 65.3602 | 65.8018 |  | B041 Cyromodule#2 | After Solenoid 1 | TX-482A | A5 | 85.3,ok
4 | 141 | X113376 | 61.476 | 60.5436 | 59.6441 | 60.0376 |  | B085 Cyromodule #9 | Before Solenoid 1 | TX-881B | A5 | 72.7 Ω, OK
| 142 | X113377 | 62.1715 | 61.2205 | 60.3032 | 61.0113 |  | B085 Cyromodule #4 | After Solenoid 3 | TX-886A | A6 | 82.0, ok
| 143 | X113378 | 63.5992 | 62.6218 | 61.679 | 61.9201 |  | B085 Cyromodule #4 | Before Solenoid 1 | TX-881B | A5 | 83.4, ok
| 144 | X113379 | 62.2145 | 61.2654 | 60.3499 | 61.4164 |  | B085 Cyromodule#7 | After Solenoid 2 | TX-884A | A5 | 80.6 Ω, OK
| 145 | X113380 | 61.1998 | 60.2731 | 59.3789 | 60.1776
| 146 | X113381 | 63.1636 | 62.2003 | 61.271 | 61.4735 |  | B085 Cyromodule #4 | After Solenoid 2 | TX-884A | A5 | 82.4, ok
| 147 | X113382 | 61.7764 | 60.8382 | 59.9331 | 61.3672 |  | B085 Cyromodule #4 | After Solenoid 1 | TX-882B | A5 | 81.2, ok
| 148 | X113383 | 63.8058 | 62.8231 | 61.8752 | 62.5472 |  | B085 Cyromodule #9 | Before Solenoid 2 | TX-883B | A5 | 62.9 Ω, OK
| 149 | X113384 | 62.9361 | 61.9692 | 61.0365 | 61.7659 |  | B085 Cyromodule #9 | After Solenoid 3 | TX-886B | A6 | 62.3 Ω, OK
| 150 | X113404 | 63.1821 | 62.2192 | 61.2901 | 62.3137
| 151 | X113405 | 63.595 | 62.623 | 61.6853 | 62.3498 |  | B085 Cyromodule #8 | Before Solenoid 1 | TX-881B | A5 | 83.4, ok
| 152 | X113436 | 64.8696 | 63.8315 | 62.8296 | 63.7276 |  | B085 Cyromodule #4 | Before Solenoid3 | TX-885B | A6 | 84.3, ok
| 153 | X113440 | 65.3366 | 64.2849 | 63.2702 | 64.0058
| 154 | X113441 | 70.3086 | 69.1636 | 68.0588 | 69.1395
| 155 | X113442 | 66.6313 | 65.5583 | 64.5229 | 65.6113
| 156 | X113443 | 69.2504 | 68.1103 | 67.0104 | 67.8192 |  | B085 Cyromodule #8 | Before Solenoid 2 | TX-883B | A5 | 87.9, ok
| 157 | X113444 | 64.3371 | 63.3043 | 62.3076 | 62.8643 |  | B085 Cyromodule #9 | After Solenoid 1 | TX-882B | A5 | 63.5 Ω, OK
| 158 | X113445 | 65.131 | 64.081 | 63.0677 | 63.7152 |  | B085 Cyromodule#7 | After Solenoid 1 | TX-882B | A5 | 84.7 Ω, OK
| 159 | X113446 | 68.4623 | 67.3509 | 66.2784 | 67.4542 |  | B085 Cyromodule #4 | Before Solenoid 2 | TX-883A | A5 | 84.3, ok
| 160 | X113448 | 67.0954 | 66.0212 | 64.9845 | 66.2605 |  | B085 Cyromodule #8 | Before Solenoid 3 | TX-885A | A6 | 85.8, ok
| 161 | X113449 | 66.9637 | 65.8808 | 64.8357 | 65.5067 |  | B085 Cyromodule #9 | After Solenoid 2 | TX-884B | A5 | 65.7 Ω, OK
| 162 | X113450 | 65.1992 | 64.1492 | 63.136 | 63.2803 |  | B085 Cyromodule #8 | Before Solenoid 2 | TX-883A | A5 | 84.2, ok
| 163 | X113452 | 66.3263 | 65.2549 | 64.2212 | 64.7253 |  | B085 Cyromodule#7 | Before Solenoid3 | TX-885B | A6 | 85.5 Ω, OK
| 164 | X113454 | 64.9077 | 63.8578 | 62.8449 | 63.7569 |  | B085 Cyromodule #8 | After Solenoid 2 | TX-884A | A5 | 83.4, ok
| 165 | X113467 | 62.6327 | 61.676 | 60.7531 | 56.0283 | Out of range - Returned to Lakeshore
| 166 | X113468 | 63.7317 | 62.7578 | 61.8183 | 62.2975 |  | B085 Cyromodule #9 | After Solenoid 3 | TX-886A | A6 | 63.0 Ω, OK
| 167 | X113469 | 62.4954 | 61.5396 | 60.6176 | 61.6925
| 168 | X113470 | 62.6293 | 61.6787 | 60.7615 | 61.3722 |  | B085 Cyromodule#7 | After Solenoid 2 | TX-884B | A5 | 81.9 Ω, OK
| 169 | X113471 | 63.0258 | 62.06 | 61.1285 | 62.1602
| 170 | X113472 | 63.4847 | 62.5097 | 61.5692 | 61.8754 |  | B085 Cyromodule #8 | After Solenoid 1 | TX-882B | A5 | 82.7, ok
| 171 | X113473 | 63.367 | 62.4094 | 61.4853 | 62.3831 |  | B085 Cyromodule #8 | After Solenoid 2 | TX-884B | A5 | 82.2, ok
| 172 | X113474 | 61.7155 | 60.7755 | 59.8685 | 60.7515 |  | B085 Cyromodule #4 | After Solenoid 2 | TX-884B | A5 | 80.9, ok
| 173 | X113475 | 62.5823 | 61.6232 | 60.6981 | 61.7204
| 174 | X113476 | 60.8485 | 59.9305 | 59.0448 | 59.4309 |  | B085 Cyromodule#7 | After Solenoid 3 | TX-886B | A6 | 80.0 Ω, OK
| 175 | X113477 | 62.3386 | 61.3873 | 60.4696 | 61.1312
| 176 | X113478 | 55.2278 | 54.4246 | 53.6485 | 54.0836 |  | B085 Cyromodule #9 | After Solenoid 1 | TX-882A | A5 | 54.6 Ω, OK
| 177 | X113479 | 57.1442 | 56.3026 | 55.4894 | 55.681 |  | B085 Cyromodule #4 | Before Solenoid 1 | TX-881A | A5 | 77.4, ok
| 178 | X113480 | 55.5082 | 54.7013 | 53.9216 | 54.6412
| 179 | X113481 | 56.0275 | 55.2089 | 54.4177 | 54.9313 |  | B085 Cyromodule #4 | Before Solenoid 2 | TX-883B | A5 | 74.7, ok
| 180 | X113482 | 53.7383 | 52.9582 | 52.2043 | 53.0392 |  | B085 Cyromodule#7 | Before Solenoid 2 | TX-883B | A5 | 73.5 Ω, OK
| 181 | X113483 | 54.0959 | 53.3084 | 52.5473 | 53.1153 |  | B085 Cyromodule #8 | After Solenoid 1 | TX-882A | A5 | 73.3, ok
| 182 | X113484 | 56.2686 | 55.4457 | 54.6505 | 55.2376 |  | B085 Cyromodule #9 | Before Solenoid 2 | TX-883A | A5 | 55.4 Ω, OK
| 183 | X113485 | 54.7698 | 53.9753 | 53.2075 | 53.3867 |  | B085 Cyromodule #4 | After Solenoid 1 | TX-882A | A5 | 77.4, ok
| 184 | X113486 | 53.5663 | 52.7927 | 52.0451 | 52.6585 |  | B085 Cyromodule #9 | Before Solenoid 3 | TX-885A | A6 | 52.9 Ω, OK
| 185 | X113487 | 56.7819 | 55.9469 | 55.1402 | 61.7084 | Out of range - Returned to Lakeshore
| 186 | X113488 | 56.1104 | 55.2895 | 54.4963 | 55.1769 |  | B085 Cyromodule #9 | Before Solenoid 1 | TX-881A | A5 | 55.4 Ω, OK
| 187 | X113489 | 56.1006 | 55.281 | 54.489 | 54.9436 |  | B085 Cyromodule#7 | Before Solenoid 3 | TX-885A | A6 | 76.1 Ω, OK
| 188 | X113490 | 55.815 | 55.0027 | 54.2178 | 54.5372 |  | B085 Cyromodule#7 | After Solenoid 3 | TX-886A | A6 | 76.1 Ω, OK
| 189 | X113491 | 55.6655 | 54.856 | 54.0736 | 54.5658 |  | B085 Cyromodule #9 | After Solenoid 2 | TX-884A | A5 | 54.8 Ω, OK
| 190 | X113492 | 55.8876 | 55.0728 | 54.2855 | 54.8882
| 191 | X113494 | 57.3026 | 56.4582 | 55.6423 | 56.2255
| 192 | X113495 | 56.1352 | 55.3134 | 54.5193 | 55.0945 |  | B29 Cyromodule #2 | After Solenoid 1 | TX-282B |  | 55.7 Ω,ok
| 193 | X113496 | 54.7739 | 53.9822 | 53.217 | 54.3467
| 194 | X113498 | 55.8104 | 54.9965 | 54.21 | 54.9632
| 195 | X113502 | 63.5098 | 62.5383 | 61.6011 | 61.6198 |  | B085 Cyromodule #8 | Before Solenoid 1 | TX-881A | A5 | 83.2, ok
| 196 | X113510 | 61.1395 | 60.2137 | 59.3207 | 59.9195
| 197 | X113511 | 62.8779 | 61.9212 | 60.9982 | 61.8564 |  | B085 Cyromodule #4 | Before Solenoid 3 | TX-885A | A6 | 82.6, ok
| 198 | X113512 | 61.952 | 61.0101 | 60.1014 | 61.0001 |  | B085 Cyromodule #4 | After Solenoid 3 | TX-886B | A6 | 81.8, ok
| 199 | X113513 | 63.7551 | 62.7742 | 61.8279 | 62.6057 |  | B085 Cyromodule#7 | Before Solenoid 2 | TX-883A | A5 | 82.8 Ω, OK
| 200 | X113514 | 62.2877 | 61.3395 | 60.4249 | 60.8682 |  | B085 Cyromodule #9 | Before Solenoid3 | TX-885B | A6 | 73.7 Ω, OK
R1 | 201 | X114090 | 53.2053 | 52.4622 | 51.7446 | 52.3796
| 202 | X114098 | 58.9677 | 58.1007 | 57.2637 | 58.0021
5 | 203 | X113375 | 61.4874 | 60.5499 | 59.6455 | 61.3735
| 204 | X113385 | 62.9028 | 61.9415 | 61.0142 | 62.3765
| 205 | X113386 | 62.6789 | 61.7228 | 60.8005 | 61.6157
| 206 | X113387 | 61.7939 | 60.8561 | 59.9516 | 61.1715
| 207 | X113500 | 61.0996 | 60.1731 | 59.2794 | 60.3423
| 208 | X113501 | 60.1858 | 59.2802 | 58.4064 | 59.4512
| 209 | X113503 | 62.7464 | 61.7911 | 60.8694 | 61.8695 |  | B085 Cyromodule #12 | After Solenoid 2 | TX-884B | A5 | 61.5
| 210 | X113504 | 63.0968 | 62.1356 | 61.2082 | 62.4778 |  | B085 Cyromodule #12 | Before Solenoid 2 | TX-883B | A5 | 62.1
| 211 | X113505 | 63.1528 | 62.184 | 61.2494 | 62.3557
| 212 | X113506 | 62.3415 | 61.4115 | 60.5139 | 61.6327
| 213 | X113507 | 61.7548 | 60.8206 | 59.9193 | 61.0916
| 214 | X113508 | 60.3496 | 59.4433 | 58.5689 | 59.8025
| 215 | X113509 | 60.3349 | 59.4326 | 58.562 | 59.8012
| 216 | X113515 | 62.5677 | 61.6095 | 60.6853 | 61.8713
| 217 | X113516 | 63.4195 | 62.45 | 61.5147 | 62.7352 |  | B29 Cyromodule #9 | After Solenoid 1 | TX-282B | A5 | 62.5
| 218 | X113517 | 63.8306 | 62.8501 | 61.9041 | 62.9581
| 219 | X113518 | 53.5857 | 52.8104 | 52.0612 | 52.9278 |  | B29 Cyromodule #9 | After Solenoid 1 | TX-282A | A5 | 52.8
| 220 | X113520 | 56.1471 | 55.3248 | 54.5303 | 55.3759 |  | B29 Cyromodule #8 | Before Solenoid 1 | TX-281A | A5 | 55.0
| 221 | X113521 | 54.803 | 54.0097 | 53.2429 | 54.0145 |  | B29 Cyromodule #9 | Before Solenoid 1 | TX-281B | A5 | 53.9
| 222 | X113522 | 57.9102 | 57.0874 | 56.2921 | 56.8453 |  | B29 Cyromodule #8 | After Solenoid 1 | TX-282B | A5 | 56.7
| 223 | X113523 | 55.3232 | 54.5217 | 53.7472 | 54.5242 |  | B29 Cyromodule #8 | Before Solenoid 1 | TX-281B | A5 | 54.2
| 224 | X113524 | 53.6002 | 52.821 | 52.0679 | 53.0821 |  | B29 Cyromodule #9 | Before Solenoid 1 | TX-281A | A5 | 52.7
| 225 | X113528 | 53.6463 | 52.8702 | 52.1202 | 52.9693 |  | B29 Cyromodule #8 | After Solenoid 1 | TX-282A | A5 | 52.5
| 226 | X113703 | 61.3029 | 60.4178 | 59.5633 | 60.6919
| 227 | X113704 | 58.6761 | 57.8121 | 56.9781 | 58.1775
| 228 | X113705 | 59.7672 | 58.8839 | 58.0312 | 58.8978
| 229 | X113706 | 58.3151 | 57.4658 | 56.6458 | 57.5562
| 230 | X113707 | 58.8959 | 58.0209 | 57.1763 | 58.6534
| 231 | X113708 | 55.0484 | 54.2585 | 53.4959 | 54.6491
| 232 | X113709 | 59.0154 | 58.1519 | 57.3183 | 58.1956
| 233 | X113710 | 58.5182 | 57.6569 | 56.8256 | 57.7464 |  | B085 Cyromodule #12 | Before Solenoid 1 | TX-881B | A5 | 57.7
| 234 | X113711 | 58.8112 | 57.9432 | 57.1052 | 58.1432
| 235 | X113718 | 60.606 | 59.6974 | 58.8204 | 59.6842
| 236 | X113719 | 59.525 | 58.6457 | 57.797 | 58.5851
| 237 | X113720 | 57.5533 | 56.7305 | 55.936 | 57.1912
| 238 | X113721 | 57.5533 | 56.7305 | 55.936 | 55.9747
| 239 | X113722 | 56.9947 | 56.1732 | 55.3802 | 56.1135
| 240 | X113723 | 56.2148 | 55.4193 | 54.6515 | 55.4734 |  | B085 Cyromodule #12 | After Solenoid 3 | TX-886A | A5 | 55.3
| 241 | X113724 | 56.5296 | 55.7092 | 54.9173 | 55.8098
| 242 | X113736 | 58.1049 | 57.2627 | 56.4496 | 57.3383
| 243 | X113737 | 55.5871 | 54.783 | 54.0069 | 55.1362
| 244 | X113738 | 55.5873 | 54.78 | 54.0008 | 54.8132 |  | B085 Cyromodule #12 | Before Solenoid 1 | TX-881A | A5 | 54.9
| 245 | X113739 | 58.6091 | 57.7504 | 56.9215 | 57.3281
| 246 | X113740 | 57.4617 | 56.6243 | 55.816 | 56.9815
| 247 | X113741 | 60.9024 | 59.9937 | 59.1167 | 60.3316
| 248 | X113742 | 56.1177 | 55.3077 | 54.5258 | 55.1145
| 249 | X113743 | 55.5341 | 54.7435 | 53.9802 | 54.5412
| 250 | X113744 | 61.6073 | 60.7054 | 59.8349 | 61.0301
| 251 | X113745 | 56.5877 | 55.7814 | 55.0029 | 56.0802
| 252 | X113746 | 55.9372 | 55.1295 | 54.3499 | 55.1632 |  | B085 Cyromodule #12 | Before Solenoid 2 | TX-883A | A5 | 55.1
| 253 | X113747 | 58.1535 | 57.3114 | 56.4984 | 57.3614 |  | B085 Cyromodule #12 | Before Solenoid 3 | TX-885B | A5 | 57.5
| 254 | X113748 | 58.8228 | 57.9631 | 57.1332 | 58.0389
| 255 | X113749 | 55.1438 | 54.3542 | 53.592 | 54.6103
| 256 | X113750 | 55.823 | 55.0189 | 54.2427 | 55.2462
| 257 | X113751 | 58.349 | 57.5281 | 56.7353 | 57.7362 |  | B085 Cyromodule #12 | Before Solenoid 3 | TX-885A | A5 | 57.5
| 258 | X113753 | 54.999 | 54.2112 | 53.4507 | 54.9241
| 259 | X113754 | 60.4369 | 59.5311 | 58.6569 | 59.9277 |  | B085 Cyromodule #12 | After Solenoid 1 | TX-882B | A5 | 59.5
| 260 | X113755 | 60.3792 | 59.4858 | 58.6234 | 60.2225
| 261 | X113756 | 57.7907 | 56.9581 | 56.1542 | 57.0642
| 262 | X113758 | 59.3247 | 58.4412 | 57.5884 | 59.3094
| 263 | X113759 | 57.4462 | 56.6016 | 55.7864 | 56.5602
| 264 | X113760 | 57.3788 | 56.5309 | 55.7126 | 56.6671
| 265 | X113761 | 58.8743 | 58.0193 | 57.1937 | 58.0921
| 266 | X113800 | 57.0862 | 56.2635 | 55.4694 | 56.3297 |  | B085 Cyromodule #12 | After Solenoid 3 | TX-886B | A5 | 56.2
| 267 | X113816 | 60.0878 | 59.1895 | 58.3227 | 59.4852 |  | B085 Cyromodule #12 | After Solenoid 1 | TX-882A | A5 | 57.7
| 268 | X113817 | 56.8899 | 56.0648 | 55.2683 | 56.5467
| 269 | X113819 | 54.476 | 53.7037 | 52.9579 | 53.8183
| 270 | X113820 | 59.8175 | 58.9231 | 58.06 | 59.2708 |  | B085 Cyromodule #12 | After Solenoid 2 | TX-884A | A5 | 58.7
| 271 | X113821 | 59.9741 | 59.0779 | 58.2129 | 59.2832
| 272 | X113822 | 53.6527 | 52.8862 | 52.1462 | 53.0324

## Network
ls1-ca01-fths-n0001 | LS1_CA01:FTHS_N0001 | facility | 10.40.201.10 | Yes | 00:50:c2:6f:43:19 | On Back “LAN/POE” | 87F001458 | N/A | LS1-003.05 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-ca01-fths-n0002 | LS1_CA01:FTHS_N0002 | facility | 10.40.201.11 | Yes | 00:50:c2:6f:43:0d | On Back “LAN/POE” | 87F001459 | N/A | LS1-003.05 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-ca02-fths-n0001 | LS1_CA02:FTHS_N0001 | facility | 10.40.201.12 | Yes | 00:50:c2:6f:43:0c | On Back “LAN/POE” | 87F001460 | N/A | LS1-003.05 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-ca02-fths-n0002 | LS1_CA02:FTHS_N0002 | facility | 10.40.201.13 | Yes | 00:50:c2:6f:43:0b | On Back “LAN/POE” | 87F001735 | N/A | LS1-003.05 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-ca03-fths-n0001 | LS1_CA03:FTHS_N0001 | facility | 10.40.201.14 | Yes | 00:50:c2:6f:43:18 | On Back “LAN/POE” | 87F001736 | N/A | LS1-003.05 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-ca03-fths-n0002 | LS1_CA03:FTHS_N0002 | facility | 10.40.201.15 | Yes | 00:50:c2:6f:43:0a | On Back “LAN/POE” | 87F001737 | N/A | LS1-003.05 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb01-fths-n0001 | LS1_CB01:FTHS_N0001 | facility | 10.40.201.16 | Yes | 00:50:c2:6f:43:1c | On Back “LAN/POE” | 87F000220 | N/A | LS1-006.02 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb01-fths-n0002 | LS1_CB01:FTHS_N0002 | facility | 10.40.201.17 | Yes | 00:50:c2:6f:43:1b | On Back “LAN/POE” | 87F001493 | N/A | LS1-006.02 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb02-fths-n0001 | LS1_CB02:FTHS_N0001 | facility | 10.40.201.18 | Yes | 00:50:c2:6f:43:10 | On Back “LAN/POE” | 87F001494 | N/A | LS1-006.02 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb03-fths-n0001 | LS1_CB03:FTHS_N0001 | facility | 10.40.201.19 | Yes | 00:50:c2:6f:43:17 | On Back “LAN/POE” | 87F001495 | N/A | LS1-006.02 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb03-fths-n0002 | LS1_CB03:FTHS_N0002 | facility | 10.40.201.20 | Yes | 00:50:c2:6f:43:0f | On Back “LAN/POE” | 87F001496 | N/A | LS1-006.02 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb04-fths-n0001 | LS1_CB04:FTHS_N0001 | facility | 10.40.201.21 | Yes | 00:50:c2:6f:43:14 | On Back “LAN/POE” | 87F001583 | N/A | LS1-011.02 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb04-fths-n0002 | LS1_CB04:FTHS_N0002 | facility | 10.40.201.22 | Yes | 00:50:c2:6f:43:05 | On Back “LAN/POE” | 87F001584 | N/A | LS1-011.02 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb05-fths-n0001 | LS1_CB05:FTHS_N0001 | facility | 10.40.201.23 | Yes | 00:50:c2:6f:42:ff | On Back “LAN/POE” | 87F001585 | N/A | LS1-011.02 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb06-fths-n0001 | LS1_CB06:FTHS_N0001 | facility | 10.40.201.24 | Yes | 00:50:c2:6f:43:03 | On Back “LAN/POE” | 87F001282 | N/A | LS1-015.01 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb06-fths-n0002 | LS1_CB06:FTHS_N0002 | facility | 10.40.201.25 | Yes | 00:50:c2:6f:43:06 | On Back “LAN/POE” | 87F001283 | N/A | LS1-015.01 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb07-fths-n0001 | LS1_CB07:FTHS_N0001 | facility | 10.40.201.26 | Yes | 00:50:c2:6f:43:00 | On Back “LAN/POE” | 87F001284 | N/A | LS1-015.01 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb08-fths-n0001 | LS1_CB08:FTHS_N0001 | facility | 10.40.201.27 | Yes | 00:50:c2:6f:43:02 | On Back “LAN/POE” | 87F001315 | N/A | LS1-017.01 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb08-fths-n0002 | LS1_CB08:FTHS_N0002 | facility | 10.40.201.28 | Yes | 00:50:c2:6f:43:01 | On Back “LAN/POE” | 87F001316 | N/A | LS1-017.01 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb09-fths-n0001 | LS1_CB09:FTHS_N0001 | facility | 10.40.201.29 | Yes | 00:50:c2:6f:43:04 | On Back “LAN/POE” | 87F001317 | N/A | LS1-017.01 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb10-fths-n0001 | LS1_CB10:FTHS_N0001 | facility | 10.40.201.30 | Yes | 00:50:c2:6f:43:11 | On Back “LAN/POE” | 87F001394 | N/A | LS1-023.01 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb10-fths-n0002 | LS1_CB10:FTHS_N0002 | facility | 10.40.201.31 | Yes | 00:50:c2:6f:43:09 | On Back “LAN/POE” | 87F001396 | N/A | LS1-023.01 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cb11-fths-n0001 | LS1_CB11:FTHS_N0001 | facility | 10.40.201.32 | Yes | 00:50:c2:6f:43:16 | On Back “LAN/POE” | 87F001397 | N/A | LS1-023.01 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cc01-fths-n0102 | LS2_CC01:FTHS_N0102 | facility |  | Yes | 00:50:c2:6f:43:0e | On Back “LAN/POE” |  | N/A | LS2-002.08 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cc03-fths-n0304 | LS2_CC03:FTHS_N0304 | facility |  | Yes | 00:50:c2:6f:43:1a | On Back “LAN/POE” |  | N/A | LS2-005.08 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cc05-fths-n0506 | LS2_CC05:FTHS_N0506 | facility |  | Yes | 00:50:c2:6f:43:07 | On Back “LAN/POE” |  | N/A | LS2-007.08 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cc07-fths-n0708 | LS2_CC07:FTHS_N0708 | facility |  | Yes | 00:50:c2:6f:43:13 | On Back “LAN/POE” |  | N/A | LS2-010.08 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cc09-fths-n0910 | LS2_CC09:FTHS_N0910 | facility |  | Yes | 00:50:c2:6f:43:15 | On Back “LAN/POE” |  | N/A | LS2-013.08 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
ls1-cc11-fths-n1112 | LS2_CC11:FTHS_N1112 | facility |  | Yes | 00:50:c2:6f:43:12 | On Back “LAN/POE” |  | N/A | LS2-015.08 | Pedro Rodriguez | ??Cyro-con 18i Temperature Monitor?
Cyro-con 18i Temperature Monitor

## .db Parsing
hostname | ch | serves this sensor
LS1_CA01:FTHS_N0001 | A | LS1:CA01:FTS:TX481A
LS1_CA01:FTHS_N0001 | B | LS1:CA01:FTS:TX481B
LS1_CA01:FTHS_N0001 | C | LS1:CA01:FTS:TX482A
LS1_CA01:FTHS_N0001 | D | LS1:CA01:FTS:TX482B
LS1_CA01:FTHS_N0001 | E
LS1_CA01:FTHS_N0001 | F
LS1_CA01:FTHS_N0001 | G
LS1_CA01:FTHS_N0001 | H
LS1_CA01:FTHS_N0002 | A | LS1:CA01:FTS:TX485A
LS1_CA01:FTHS_N0002 | B | LS1:CA01:FTS:TX485B
LS1_CA01:FTHS_N0002 | C | LS1:CA01:FTS:TX486A
LS1_CA01:FTHS_N0002 | D | LS1:CA01:FTS:TX486B
LS1_CA01:FTHS_N0002 | E
LS1_CA01:FTHS_N0002 | F
LS1_CA01:FTHS_N0002 | G
LS1_CA01:FTHS_N0002 | H
LS1_CA02:FTHS_N0001 | A | LS1:CA02:FTS:TX481A
LS1_CA02:FTHS_N0001 | B | LS1:CA02:FTS:TX481B
LS1_CA02:FTHS_N0001 | C | LS1:CA02:FTS:TX482A
LS1_CA02:FTHS_N0001 | D | LS1:CA02:FTS:TX482B
LS1_CA02:FTHS_N0001 | E
LS1_CA02:FTHS_N0001 | F
LS1_CA02:FTHS_N0001 | G
LS1_CA02:FTHS_N0001 | H
LS1_CA02:FTHS_N0002 | A | LS1:CA02:FTS:TX485A
LS1_CA02:FTHS_N0002 | B | LS1:CA02:FTS:TX485B
LS1_CA02:FTHS_N0002 | C | LS1:CA02:FTS:TX486A
LS1_CA02:FTHS_N0002 | D | LS1:CA02:FTS:TX486B
LS1_CA02:FTHS_N0002 | E
LS1_CA02:FTHS_N0002 | F
LS1_CA02:FTHS_N0002 | G
LS1_CA02:FTHS_N0002 | H
LS1_CA03:FTHS_N0001 | A | LS1:CA03:FTS:TX481A
LS1_CA03:FTHS_N0001 | B | LS1:CA03:FTS:TX481B
LS1_CA03:FTHS_N0001 | C | LS1:CA03:FTS:TX482A
LS1_CA03:FTHS_N0001 | D | LS1:CA03:FTS:TX482B
LS1_CA03:FTHS_N0001 | E
LS1_CA03:FTHS_N0001 | F
LS1_CA03:FTHS_N0001 | G
LS1_CA03:FTHS_N0001 | H
LS1_CA03:FTHS_N0002 | A | LS1:CA03:FTS:TX485A
LS1_CA03:FTHS_N0002 | B | LS1:CA03:FTS:TX485B
LS1_CA03:FTHS_N0002 | C | LS1:CA03:FTS:TX486A
LS1_CA03:FTHS_N0002 | D | LS1:CA03:FTS:TX486B
LS1_CA03:FTHS_N0002 | E
LS1_CA03:FTHS_N0002 | F
LS1_CA03:FTHS_N0002 | G
LS1_CA03:FTHS_N0002 | H
LS1_CB01:FTHS_N0001 | A | LS1:CB01:FTHS:TX881A
LS1_CB01:FTHS_N0001 | B | LS1:CB01:FTHS:TX881B
LS1_CB01:FTHS_N0001 | C | LS1:CB01:FTHS:TX882A
LS1_CB01:FTHS_N0001 | D | LS1:CB01:FTHS:TX882B
LS1_CB01:FTHS_N0001 | E | LS1:CB01:FTHS:TX883A
LS1_CB01:FTHS_N0001 | F | LS1:CB01:FTHS:TX883B
LS1_CB01:FTHS_N0001 | G | LS1:CB01:FTHS:TX884A
LS1_CB01:FTHS_N0001 | H | LS1:CB01:FTHS:TX884B
LS1_CB01:FTHS_N0002 | A | LS1:CB01:FTHS:TX885A
LS1_CB01:FTHS_N0002 | B | LS1:CB01:FTHS:TX885B
LS1_CB01:FTHS_N0002 | C | LS1:CB01:FTHS:TX886A
LS1_CB01:FTHS_N0002 | D | LS1:CB01:FTHS:TX886A
LS1_CB01:FTHS_N0002 | E | LS1:CB02:FTHS:TX881A
LS1_CB01:FTHS_N0002 | F | LS1:CB02:FTHS:TX881B
LS1_CB01:FTHS_N0002 | G | LS1:CB02:FTHS:TX882A
LS1_CB01:FTHS_N0002 | H | LS1:CB02:FTHS:TX882B
LS1_CB02:FTHS_N0001 | A | LS1:CB02:FTHS:TX883A
LS1_CB02:FTHS_N0001 | B | LS1:CB02:FTHS:TX883B
LS1_CB02:FTHS_N0001 | C | LS1:CB02:FTHS:TX884A
LS1_CB02:FTHS_N0001 | D | LS1:CB02:FTHS:TX884B
LS1_CB02:FTHS_N0001 | E | LS1:CB02:FTHS:TX885A
LS1_CB02:FTHS_N0001 | F | LS1:CB02:FTHS:TX885B
LS1_CB02:FTHS_N0001 | G | LS1:CB02:FTHS:TX886A
LS1_CB02:FTHS_N0001 | H | LS1:CB02:FTHS:TX886A
LS1_CB03:FTHS_N0001 | A | LS1:CB03:FTHS:TX881A
LS1_CB03:FTHS_N0001 | B | LS1:CB03:FTHS:TX881B
LS1_CB03:FTHS_N0001 | C | LS1:CB03:FTHS:TX882A
LS1_CB03:FTHS_N0001 | D | LS1:CB03:FTHS:TX882B
LS1_CB03:FTHS_N0001 | E | LS1:CB03:FTHS:TX883A
LS1_CB03:FTHS_N0001 | F | LS1:CB03:FTHS:TX883B
LS1_CB03:FTHS_N0001 | G | LS1:CB03:FTHS:TX884A
LS1_CB03:FTHS_N0001 | H | LS1:CB03:FTHS:TX884B
LS1_CB03:FTHS_N0002 | A | LS1:CB03:FTHS:TX885A
LS1_CB03:FTHS_N0002 | B | LS1:CB03:FTHS:TX885B
LS1_CB03:FTHS_N0002 | C | LS1:CB03:FTHS:TX886A
LS1_CB03:FTHS_N0002 | D | LS1:CB03:FTHS:TX886A
LS1_CB03:FTHS_N0002 | E
LS1_CB03:FTHS_N0002 | F
LS1_CB03:FTHS_N0002 | G
LS1_CB03:FTHS_N0002 | H
LS1_CB04:FTHS_N0001 | A | LS1:CB04:FTHS:TX881A
LS1_CB04:FTHS_N0001 | B | LS1:CB04:FTHS:TX881B
LS1_CB04:FTHS_N0001 | C | LS1:CB04:FTHS:TX882A
LS1_CB04:FTHS_N0001 | D | LS1:CB04:FTHS:TX882B
LS1_CB04:FTHS_N0001 | E | LS1:CB04:FTHS:TX883A
LS1_CB04:FTHS_N0001 | F | LS1:CB04:FTHS:TX883B
LS1_CB04:FTHS_N0001 | G | LS1:CB04:FTHS:TX884A
LS1_CB04:FTHS_N0001 | H | LS1:CB04:FTHS:TX884B
LS1_CB04:FTHS_N0002 | A | LS1:CB04:FTHS:TX885A
LS1_CB04:FTHS_N0002 | B | LS1:CB04:FTHS:TX885B
LS1_CB04:FTHS_N0002 | C | LS1:CB04:FTHS:TX886A
LS1_CB04:FTHS_N0002 | D | LS1:CB04:FTHS:TX886A
LS1_CB04:FTHS_N0002 | E | LS1:CB05:FTHS:TX881A
LS1_CB04:FTHS_N0002 | F | LS1:CB05:FTHS:TX881B
LS1_CB04:FTHS_N0002 | G | LS1:CB05:FTHS:TX882A
LS1_CB04:FTHS_N0002 | H | LS1:CB05:FTHS:TX882B
LS1_CB05:FTHS_N0001 | A | LS1:CB05:FTHS:TX883A
LS1_CB05:FTHS_N0001 | B | LS1:CB05:FTHS:TX883B
LS1_CB05:FTHS_N0001 | C | LS1:CB05:FTHS:TX884A
LS1_CB05:FTHS_N0001 | D | LS1:CB05:FTHS:TX884B
LS1_CB05:FTHS_N0001 | E | LS1:CB05:FTHS:TX885A
LS1_CB05:FTHS_N0001 | F | LS1:CB05:FTHS:TX885B
LS1_CB05:FTHS_N0001 | G | LS1:CB05:FTHS:TX886A
LS1_CB05:FTHS_N0001 | H | LS1:CB05:FTHS:TX886A
LS1_CB06:FTHS_N0001 | A | LS1:CB06:FTHS:TX881A
LS1_CB06:FTHS_N0001 | B | LS1:CB06:FTHS:TX881B
LS1_CB06:FTHS_N0001 | C | LS1:CB06:FTHS:TX882A
LS1_CB06:FTHS_N0001 | D | LS1:CB06:FTHS:TX882B
LS1_CB06:FTHS_N0001 | E | LS1:CB06:FTHS:TX883A
LS1_CB06:FTHS_N0001 | F | LS1:CB06:FTHS:TX883B
LS1_CB06:FTHS_N0001 | G | LS1:CB06:FTHS:TX884A
LS1_CB06:FTHS_N0001 | H | LS1:CB06:FTHS:TX884B
LS1_CB06:FTHS_N0002 | A | LS1:CB06:FTHS:TX885A
LS1_CB06:FTHS_N0002 | B | LS1:CB06:FTHS:TX885B
LS1_CB06:FTHS_N0002 | C | LS1:CB06:FTHS:TX886A
LS1_CB06:FTHS_N0002 | D | LS1:CB06:FTHS:TX886A
LS1_CB06:FTHS_N0002 | E | LS1:CB07:FTHS:TX881A
LS1_CB06:FTHS_N0002 | F | LS1:CB07:FTHS:TX881B
LS1_CB06:FTHS_N0002 | G | LS1:CB07:FTHS:TX882A
LS1_CB06:FTHS_N0002 | H | LS1:CB07:FTHS:TX882B
LS1_CB07:FTHS_N0001 | A | LS1:CB07:FTHS:TX883A
LS1_CB07:FTHS_N0001 | B | LS1:CB07:FTHS:TX883B
LS1_CB07:FTHS_N0001 | C | LS1:CB07:FTHS:TX884A
LS1_CB07:FTHS_N0001 | D | LS1:CB07:FTHS:TX884B
LS1_CB07:FTHS_N0001 | E | LS1:CB07:FTHS:TX885A
LS1_CB07:FTHS_N0001 | F | LS1:CB07:FTHS:TX885B
LS1_CB07:FTHS_N0001 | G | LS1:CB07:FTHS:TX886A
LS1_CB07:FTHS_N0001 | H | LS1:CB07:FTHS:TX886A
LS1_CB08:FTHS_N0001 | A | LS1:CB08:FTHS:TX881A
LS1_CB08:FTHS_N0001 | B | LS1:CB08:FTHS:TX881B
LS1_CB08:FTHS_N0001 | C | LS1:CB08:FTHS:TX882A
LS1_CB08:FTHS_N0001 | D | LS1:CB08:FTHS:TX882B
LS1_CB08:FTHS_N0001 | E | LS1:CB08:FTHS:TX883A
LS1_CB08:FTHS_N0001 | F | LS1:CB08:FTHS:TX883B
LS1_CB08:FTHS_N0001 | G | LS1:CB08:FTHS:TX884A
LS1_CB08:FTHS_N0001 | H | LS1:CB08:FTHS:TX884B
LS1_CB08:FTHS_N0002 | A | LS1:CB08:FTHS:TX885A
LS1_CB08:FTHS_N0002 | B | LS1:CB08:FTHS:TX885B
LS1_CB08:FTHS_N0002 | C | LS1:CB08:FTHS:TX886A
LS1_CB08:FTHS_N0002 | D | LS1:CB08:FTHS:TX886A
LS1_CB08:FTHS_N0002 | E | LS1:CB09:FTHS:TX881A
LS1_CB08:FTHS_N0002 | F | LS1:CB09:FTHS:TX881B
LS1_CB08:FTHS_N0002 | G | LS1:CB09:FTHS:TX882A
LS1_CB08:FTHS_N0002 | H | LS1:CB09:FTHS:TX882B
LS1_CB09:FTHS_N0001 | A | LS1:CB09:FTHS:TX883A
LS1_CB09:FTHS_N0001 | B | LS1:CB09:FTHS:TX883B
LS1_CB09:FTHS_N0001 | C | LS1:CB09:FTHS:TX884A
LS1_CB09:FTHS_N0001 | D | LS1:CB09:FTHS:TX884B
LS1_CB09:FTHS_N0001 | E | LS1:CB09:FTHS:TX885A
LS1_CB09:FTHS_N0001 | F | LS1:CB09:FTHS:TX885B
LS1_CB09:FTHS_N0001 | G | LS1:CB09:FTHS:TX886A
LS1_CB09:FTHS_N0001 | H | LS1:CB09:FTHS:TX886A
LS1_CB10:FTHS_N0001 | A | LS1:CB10:FTHS:TX881A
LS1_CB10:FTHS_N0001 | B | LS1:CB10:FTHS:TX881B
LS1_CB10:FTHS_N0001 | C | LS1:CB10:FTHS:TX882A
LS1_CB10:FTHS_N0001 | D | LS1:CB10:FTHS:TX882B
LS1_CB10:FTHS_N0001 | E | LS1:CB10:FTHS:TX883A
LS1_CB10:FTHS_N0001 | F | LS1:CB10:FTHS:TX883B
LS1_CB10:FTHS_N0001 | G | LS1:CB10:FTHS:TX884A
LS1_CB10:FTHS_N0001 | H | LS1:CB10:FTHS:TX884B
LS1_CB10:FTHS_N0002 | A | LS1:CB10:FTHS:TX885A
LS1_CB10:FTHS_N0002 | B | LS1:CB10:FTHS:TX885B
LS1_CB10:FTHS_N0002 | C | LS1:CB10:FTHS:TX886A
LS1_CB10:FTHS_N0002 | D | LS1:CB10:FTHS:TX886A
LS1_CB10:FTHS_N0002 | E | LS1:CB11:FTHS:TX881A
LS1_CB10:FTHS_N0002 | F | LS1:CB11:FTHS:TX881B
LS1_CB10:FTHS_N0002 | G | LS1:CB11:FTHS:TX882A
LS1_CB10:FTHS_N0002 | H | LS1:CB11:FTHS:TX882B
LS1_CB11:FTHS_N0001 | A | LS1:CB11:FTHS:TX883A
LS1_CB11:FTHS_N0001 | B | LS1:CB11:FTHS:TX883B
LS1_CB11:FTHS_N0001 | C | LS1:CB11:FTHS:TX884A
LS1_CB11:FTHS_N0001 | D | LS1:CB11:FTHS:TX884B
LS1_CB11:FTHS_N0001 | E | LS1:CB11:FTHS:TX885A
LS1_CB11:FTHS_N0001 | F | LS1:CB11:FTHS:TX885B
LS1_CB11:FTHS_N0001 | G | LS1:CB11:FTHS:TX886A
LS1_CB11:FTHS_N0001 | H | LS1:CB11:FTHS:TX886A
LS2_CC01:FTHS_N0102 | A | LS1:CC01:FTS:TX281A
LS2_CC01:FTHS_N0102 | B | LS1:CC01:FTS:TX281B
LS2_CC01:FTHS_N0102 | C | LS1:CC01:FTS:TX282A
LS2_CC01:FTHS_N0102 | D | LS1:CC01:FTS:TX282B
LS2_CC01:FTHS_N0102 | E | LS1:CC02:FTS:TX281A
LS2_CC01:FTHS_N0102 | F | LS1:CC02:FTS:TX281B
LS2_CC01:FTHS_N0102 | G | LS1:CC02:FTS:TX282A
LS2_CC01:FTHS_N0102 | H | LS1:CC02:FTS:TX282B
LS2_CC03:FTHS_N0304 | A | LS1:CC03:FTS:TX281A
LS2_CC03:FTHS_N0304 | B | LS1:CC03:FTS:TX281B
LS2_CC03:FTHS_N0304 | C | LS1:CC03:FTS:TX282A
LS2_CC03:FTHS_N0304 | D | LS1:CC03:FTS:TX282B
LS2_CC03:FTHS_N0304 | E | LS1:CC04:FTS:TX281A
LS2_CC03:FTHS_N0304 | F | LS1:CC04:FTS:TX281B
LS2_CC03:FTHS_N0304 | G | LS1:CC04:FTS:TX282A
LS2_CC03:FTHS_N0304 | H | LS1:CC04:FTS:TX282B
LS2_CC05:FTHS_N0506 | A | LS1:CC05:FTS:TX281A
LS2_CC05:FTHS_N0506 | B | LS1:CC05:FTS:TX281B
LS2_CC05:FTHS_N0506 | C | LS1:CC05:FTS:TX282A
LS2_CC05:FTHS_N0506 | D | LS1:CC05:FTS:TX282B
LS2_CC05:FTHS_N0506 | E | LS1:CC06:FTS:TX281A
LS2_CC05:FTHS_N0506 | F | LS1:CC06:FTS:TX281B
LS2_CC05:FTHS_N0506 | G | LS1:CC06:FTS:TX282A
LS2_CC05:FTHS_N0506 | H | LS1:CC06:FTS:TX282B
LS2_CC07:FTHS_N0708 | A | LS1:CC07:FTS:TX281A
LS2_CC07:FTHS_N0708 | B | LS1:CC07:FTS:TX281B
LS2_CC07:FTHS_N0708 | C | LS1:CC07:FTS:TX282A
LS2_CC07:FTHS_N0708 | D | LS1:CC07:FTS:TX282B
LS2_CC07:FTHS_N0708 | E | LS1:CC08:FTS:TX281A
LS2_CC07:FTHS_N0708 | F | LS1:CC08:FTS:TX281B
LS2_CC07:FTHS_N0708 | G | LS1:CC08:FTS:TX282A
LS2_CC07:FTHS_N0708 | H | LS1:CC08:FTS:TX282B
LS2_CC09:FTHS_N0910 | A | LS1:CC09:FTS:TX281A
LS2_CC09:FTHS_N0910 | B | LS1:CC09:FTS:TX281B
LS2_CC09:FTHS_N0910 | C | LS1:CC09:FTS:TX282A
LS2_CC09:FTHS_N0910 | D | LS1:CC09:FTS:TX282B
LS2_CC09:FTHS_N0910 | E | LS1:CC10:FTS:TX281A
LS2_CC09:FTHS_N0910 | F | LS1:CC10:FTS:TX281B
LS2_CC09:FTHS_N0910 | G | LS1:CC10:FTS:TX282A
LS2_CC09:FTHS_N0910 | H | LS1:CC10:FTS:TX282B
LS2_CC11:FTHS_N1112 | A | LS1:CC11:FTS:TX281A
LS2_CC11:FTHS_N1112 | B | LS1:CC11:FTS:TX281B
LS2_CC11:FTHS_N1112 | C | LS1:CC11:FTS:TX282A
LS2_CC11:FTHS_N1112 | D | LS1:CC11:FTS:TX282B
LS2_CC11:FTHS_N1112 | E | LS1:CC12:FTS:TX281A
LS2_CC11:FTHS_N1112 | F | LS1:CC12:FTS:TX281B
LS2_CC11:FTHS_N1112 | G | LS1:CC12:FTS:TX282A
LS2_CC11:FTHS_N1112 | H | LS1:CC12:FTS:TX282B

## CM-Sensor to IOC routing
CM Position | CM build | sensor position | Sensor SN | R@room temp | 1010/ 1050 | Cable number | wiring group* | Rack | Device Name | Device Input Channel | Output PV
CB01 | SCM811 | TX-881A | x111742 | 85.231 | 1050 | 34E000574 | A | LS1-N0602 | LS1-CB01 | A
CB01 | SCM811 | TX-881B | x111773 | 85.2763 | 1050 | 34E000574 | B | LS1-N0602 | LS1-CB01 | B
CB01 | SCM811 | TX-882A | x111720 | 84.382 | 1050 | 34E000574 | C | LS1-N0602 | LS1-CB01 | C
CB01 | SCM811 | TX-882B | x111781 | 85.6811 | 1050 | 34E000574 | D | LS1-N0602 | LS1-CB01 | D
CB01 | SCM811 | TX-883A | x111714 | 85.6069 | 1050 | 34E000582 | A | LS1-N0602 | LS1-CB01 | E
CB01 | SCM811 | TX-883B | x111679 | 83.7585 | 1050 | 34E000582 | B | LS1-N0602 | LS1-CB01 | F
CB01 | SCM811 | TX-884A | x111684 | 84.5424 | 1050 | 34E000582 | C | LS1-N0602 | LS1-CB01 | G
CB01 | SCM811 | TX-884B | x111689 | 85.9847 | 1050 | 34E000582 | D | LS1-N0602 | LS1-CB01 | H
CB01 | SCM811 | TX-885A | x111678 | 85.8764 | 1050 | 34E000581 | A | LS1-N0602 | LS1-CB01-CB02 | A
CB01 | SCM811 | TX-885B | x111671 | 84.0036 | 1050 | 34E000581 | B | LS1-N0602 | LS1-CB01-CB02 | B
CB01 | SCM811 | TX-886A | x111646 | 84.9751 | 1050 | 34E000581 | C | LS1-N0602 | LS1-CB01-CB02 | C
CB01 | SCM811 | TX-886B | x111674 | 83.4126 | 1050 | 34E000581 | D | LS1-N0602 | LS1-CB01-CB02 | D
CB02 | SCM809 | TX-881A | x113488 | 55.2895 | 1050 | 34E000580 | A | LS1-N0602 | LS1-CB01-CB02 | E
CB02 | SCM809 | TX-881B | x113376 | 60.5436 | 1050 | 34E000580 | B | LS1-N0602 | LS1-CB01-CB02 | F
CB02 | SCM809 | TX-882A | x113478 | 54.4246 | 1050 | 34E000580 | C | LS1-N0602 | LS1-CB01-CB02 | G
CB02 | SCM809 | TX-882B | x113444 | 63.3043 | 1050 | 34E000580 | D | LS1-N0602 | LS1-CB01-CB02 | H
CB02 | SCM809 | TX-883A | x113484 | 55.4457 | 1050 | 34E000576 | A | LS1-N0602 | LS1-CB02 | A
CB02 | SCM809 | TX-883B | x113383 | 62.8231 | 1050 | 34E000576 | B | LS1-N0602 | LS1-CB02 | B
CB02 | SCM809 | TX-884A | x113491 | 54.856 | 1050 | 34E000576 | C | LS1-N0602 | LS1-CB02 | C
CB02 | SCM809 | TX-884B | x113449 | 65.8808 | 1050 | 34E000576 | D | LS1-N0602 | LS1-CB02 | D
CB02 | SCM809 | TX-885A | x113486 | 52.7927 | 1050 | 34E000579 | A | LS1-N0602 | LS1-CB02 | E
CB02 | SCM809 | TX-885B | x113514 | 61.3395 | 1050 | 34E000579 | B | LS1-N0602 | LS1-CB02 | F
CB02 | SCM809 | TX-886A | x113468 | 62.7578 | 1050 | 34E000579 | C | LS1-N0602 | LS1-CB02 | G
CB02 | SCM809 | TX-886B | x113384 | 61.9692 | 1050 | 34E000579 | D | LS1-N0602 | LS1-CB02 | H
CB03 | SCM810 | TX-881A | x111769 | 84.6108 | 1050 | 34E000578 | A | LS1-N0602 | LS1-CB03-1 | A
CB03 | SCM810 | TX-881B | x111782 | 83.6988 | 1050 | 34E000578 | B | LS1-N0602 | LS1-CB03-1 | B
CB03 | SCM810 | TX-882A | x111643 | 85.5847 | 1050 | 34E000578 | C | LS1-N0602 | LS1-CB03-1 | C
CB03 | SCM810 | TX-882B | x111668 | 83.889 | 1050 | 34E000578 | D | LS1-N0602 | LS1-CB03-1 | D
CB03 | SCM810 | TX-883A | x111722 | 85.7062 | 1050 | 34E000577 | A | LS1-N0602 | LS1-CB03-1 | E
CB03 | SCM810 | TX-883B | x111716 | 85.0114 | 1050 | 34E000577 | B | LS1-N0602 | LS1-CB03-1 | F
CB03 | SCM810 | TX-884A | x111758 | 85.1409 | 1050 | 34E000577 | C | LS1-N0602 | LS1-CB03-1 | G
CB03 | SCM810 | TX-884B | x111789 | 83.9901 | 1050 | 34E000577 | D | LS1-N0602 | LS1-CB03-1 | H
CB03 | SCM810 | TX-885A | x111771 | 85.8647 | 1050 | 34E000575 | A | LS1-N0602 | LS1-CB03-2 | A
CB03 | SCM810 | TX-885B | x111778 | 83.9921 | 1050 | 34E000575 | B | LS1-N0602 | LS1-CB03-2 | B
CB03 | SCM810 | TX-886A | x111644 | 84.075 | 1050 | 34E000575 | C | LS1-N0602 | LS1-CB03-2 | C
CB03 | SCM810 | TX-886B | x111721 | 83.8034 | 1050 | 34E000575 | D | LS1-N0602 | LS1-CB03-2 | D
CB04 | SCM808 | TX-881A | x113502 | 62.5383 | 1050 | 34E000584 | A | LS1-N1102 | LS1-CB04 | A
CB04 | SCM808 | TX-881B | x113405 | 62.62 | 1050 | 34E000584 | B | LS1-N1102 | LS1-CB04 | B
CB04 | SCM808 | TX-882A | x113483 | 53.3084 | 1050 | 34E000584 | C | LS1-N1102 | LS1-CB04 | C
CB04 | SCM808 | TX-882B | x113472 | 62.5 | 1050 | 34E000584 | D | LS1-N1102 | LS1-CB04 | D
CB04 | SCM808 | TX-883A | x113450 | 64.1492 | 1050 | 34E000587 | A | LS1-N1102 | LS1-CB04 | E
CB04 | SCM808 | TX-883B | x113443 | 68.1103 | 1050 | 34E000587 | B | LS1-N1102 | LS1-CB04 | F
CB04 | SCM808 | TX-884A | x113454 | 63.8578 | 1050 | 34E000587 | C | LS1-N1102 | LS1-CB04 | G
CB04 | SCM808 | TX-884B | x113473 | 62.4094 | 1050 | 34E000587 | D | LS1-N1102 | LS1-CB04 | H
CB04 | SCM808 | TX-885A | x113448 | 66.0212 | 1050 | 34E000588 | A | LS1-N1102 | LS1-CB04-CB05 | A
CB04 | SCM808 | TX-885B | x112415 | 66.3334 | 1050 | 34E000588 | B | LS1-N1102 | LS1-CB04-CB05 | B
CB04 | SCM808 | TX-886A | x112356 | 78.1917 | 1050 | 34E000588 | C | LS1-N1102 | LS1-CB04-CB05 | C
CB04 | SCM808 | TX-886B | x112403 | 80.9873 | 1050 | 34E000588 | D | LS1-N1102 | LS1-CB04-CB05 | D
CB05 | SCM807 | TX-881A | x112425 | 74.9701 | 1050 | 34E000585 | A | LS1-N1102 | LS1-CB04-CB05 | E
CB05 | SCM807 | TX-881B | x112426 | 74.1722 | 1050 | 34E000585 | B | LS1-N1102 | LS1-CB04-CB05 | F
CB05 | SCM807 | TX-882A | x112424 | 70.5614 | 1050 | 34E000585 | C | LS1-N1102 | LS1-CB04-CB05 | G
CB05 | SCM807 | TX-882B | x113445 | 64.081 | 1050 | 34E000585 | D | LS1-N1102 | LS1-CB04-CB05 | H
CB05 | SCM807 | TX-883A | x113513 | 62.7742 | 1050 | 34E000583 | A | LS1-N1102 | LS1-CB05 | A
CB05 | SCM807 | TX-883B | x113482 | 52.9582 | 1050 | 34E000583 | B | LS1-N1102 | LS1-CB05 | B
CB05 | SCM807 | TX-884A | x113379 | 61.2654 | 1050 | 34E000583 | C | LS1-N1102 | LS1-CB05 | C
CB05 | SCM807 | TX-884B | x113470 | 61.6787 | 1050 | 34E000583 | D | LS1-N1102 | LS1-CB05 | D
CB05 | SCM807 | TX-885A | x113489 | 55.281 | 1050 | 34E000586 | A | LS1-N1102 | LS1-CB05 | E
CB05 | SCM807 | TX-885B | x113452 | 65.2549 | 1050 | 34E000586 | B | LS1-N1102 | LS1-CB05 | F
CB05 | SCM807 | TX-886A | x113490 | 55.0027 | 1050 | 34E000586 | C | LS1-N1102 | LS1-CB05 | G
CB05 | SCM807 | TX-886B | x113476 | 59.9305 | 1050 | 34E000586 | D | LS1-N1102 | LS1-CB05 | H
CB06 | SCM805 | TX-881A | x113441 | 69.16 | 1050 | 34E000589 | A | LS1-N1501 | LS1-CB06 | A
CB06 | SCM805 | TX-881B | x113477 | 61.39 | 1050 | 34E000589 | B | LS1-N1501 | LS1-CB06 | B
CB06 | SCM805 | TX-882A | x112420 | 70.44 | 1050 | 34E000589 | C | LS1-N1501 | LS1-CB06 | C
CB06 | SCM805 | TX-882B | x113380 | 60.27 | 1050 | 34E000589 | D | LS1-N1501 | LS1-CB06 | D
CB06 | SCM805 | TX-883A | x113510 | 60.21 | 1050 | 34E000591 | A | LS1-N1501 | LS1-CB06 | E
CB06 | SCM805 | TX-883B | x113440 | 64.28 | 1050 | 34E000591 | B | LS1-N1501 | LS1-CB06 | F
CB06 | SCM805 | TX-884A | x113442 | 65.55 | 1050 | 34E000591 | C | LS1-N1501 | LS1-CB06 | G
CB06 | SCM805 | TX-884B | x113494 | 56.45 | 1050 | 34E000591 | D | LS1-N1501 | LS1-CB06 | H
CB06 | SCM805 | TX-885A | x113492 | 55.07 | 1050 | 34E000593 | A | LS1-N1501 | LS1-CB06-CB07 | A
CB06 | SCM805 | TX-885B | x113498 | 54.99 | 1050 | 34E000593 | B | LS1-N1501 | LS1-CB06-CB07 | B
CB06 | SCM805 | TX-886A | x113496 | 53.98 | 1050 | 34E000593 | C | LS1-N1501 | LS1-CB06-CB07 | C
CB06 | SCM805 | TX-886B | x113475 | 61.62 | 1050 | 34E000593 | D | LS1-N1501 | LS1-CB06-CB07 | D
CB07 | SCM806 | TX-881A | x112348 | 76.7115 | 1050 | 34E000594 | A | LS1-N1501 | LS1-CB06-CB07 | E
CB07 | SCM806 | TX-881B | x112347 | 74.4938 | 1050 | 34E000594 | B | LS1-N1501 | LS1-CB06-CB07 | F
CB07 | SCM806 | TX-882A | x112346 | 70.711 | 1050 | 34E000594 | C | LS1-N1501 | LS1-CB06-CB07 | G
CB07 | SCM806 | TX-882B | x112276 | 73.5645 | 1050 | 34E000594 | D | LS1-N1501 | LS1-CB06-CB07 | H
CB07 | SCM806 | TX-883A | x112350 | 68.9416 | 1050 | 34E000590 | A | LS1-N1501 | LS1-CB07 | A
CB07 | SCM806 | TX-883B | x110679 | 85.8282 | 1050 | 34E000590 | B | LS1-N1501 | LS1-CB07 | B
CB07 | SCM806 | TX-884A | x112343 | 77.7049 | 1050 | 34E000590 | C | LS1-N1501 | LS1-CB07 | C
CB07 | SCM806 | TX-884B | x112351 | 78.363 | 1050 | 34E000590 | D | LS1-N1501 | LS1-CB07 | D
CB07 | SCM806 | TX-885A | x112349 | 74.5603 | 1050 | 34E000592 | A | LS1-N1501 | LS1-CB07 | E
CB07 | SCM806 | TX-885B | x112345 | 78.8203 | 1050 | 34E000592 | B | LS1-N1501 | LS1-CB07 | F
CB07 | SCM806 | TX-886A | x112422 | 74.7496 | 1050 | 34E000592 | C | LS1-N1501 | LS1-CB07 | G
CB07 | SCM806 | TX-886B | x112421 | 79.02 | 1050 | 34E000592 | D | LS1-N1501 | LS1-CB07 | H
CB08 | SCM804 | TX-881A | x113479 | 56.3 | 1050 | 34E000595 | A | LS1-N1701 | LS1-CB08 | A
CB08 | SCM804 | TX-881B | x113378 | 62.6 | 1050 | 34E000595 | B | LS1-N1701 | LS1-CB08 | B
CB08 | SCM804 | TX-882A | x113485 | 54 | 1050 | 34E000595 | C | LS1-N1701 | LS1-CB08 | C
CB08 | SCM804 | TX-882B | x113382 | 60.5 | 1050 | 34E000595 | D | LS1-N1701 | LS1-CB08 | D
CB08 | SCM804 | TX-883A | x113446 | 67.4 | 1050 | 34E000598 | A | LS1-N1701 | LS1-CB08 | E
CB08 | SCM804 | TX-883B | x113481 | 55.2 | 1050 | 34E000598 | B | LS1-N1701 | LS1-CB08 | F
CB08 | SCM804 | TX-884A | x113381 | 62.2 | 1050 | 34E000598 | C | LS1-N1701 | LS1-CB08 | G
CB08 | SCM804 | TX-884B | x113474 | 60.8 | 1050 | 34E000598 | D | LS1-N1701 | LS1-CB08 | H
CB08 | SCM804 | TX-885A | x113511 | 61.9 | 1050 | 34E000597 | A | LS1-N1701 | LS1-CB08-CB09 | A
CB08 | SCM804 | TX-885B | x113436 | 63.8 | 1050 | 34E000597 | B | LS1-N1701 | LS1-CB08-CB09 | B
CB08 | SCM804 | TX-886A | x113377 | 61.2 | 1050 | 34E000597 | C | LS1-N1701 | LS1-CB08-CB09 | C
CB08 | SCM804 | TX-886B | x113512 | 61 | 1050 | 34E000597 | D | LS1-N1701 | LS1-CB08-CB09 | D
CB09 | SCM801 | TX-881A | X105264 | 34.872 | 1010 | 34E000600 | A | LS1-N1701 | LS1-CB08-CB09 | E
CB09 | SCM801 | TX-881B | X105214 | 34.122 | 1010 | 34E000600 | B | LS1-N1701 | LS1-CB08-CB09 | F
CB09 | SCM801 | TX-882A | X105288 | 35.76 | 1010 | 34E000600 | C | LS1-N1701 | LS1-CB08-CB09 | G
CB09 | SCM801 | TX-882B | X105252 | 35.49 | 1010 | 34E000600 | D | LS1-N1701 | LS1-CB08-CB09 | H
CB09 | SCM801 | TX-883A | X105229 | 34.14 | 1010 | 34E000599 | A | LS1-N1701 | LS1-CB09 | A
CB09 | SCM801 | TX-883B | X105211 | 35.373 | 1010 | 34E000599 | B | LS1-N1701 | LS1-CB09 | B
CB09 | SCM801 | TX-884A | X105275 | 35.697 | 1010 | 34E000599 | C | LS1-N1701 | LS1-CB09 | C
CB09 | SCM801 | TX-884B | X105251 | 34.56 | 1010 | 34E000599 | D | LS1-N1701 | LS1-CB09 | D
CB09 | SCM801 | TX-885A | X105286 | 34.325 | 1010 | 34E000596 | A | LS1-N1701 | LS1-CB09 | E
CB09 | SCM801 | TX-885B | X105287 | 35.741 | 1010 | 34E000596 | B | LS1-N1701 | LS1-CB09 | F
CB09 | SCM801 | TX-886A | X105204 | 34.881 | 1010 | 34E000596 | C | LS1-N1701 | LS1-CB09 | G
CB09 | SCM801 | TX-886B | X105218 | 35.138 | 1010 | 34E000596 | D | LS1-N1701 | LS1-CB09 | H
CB10 | SCM802 | TX-881A | x110458 | 29.884 | 1010 | 34E000601 | A | LS1-N2301 | LS1-CB10 | A
CB10 | SCM802 | TX-881B | x110529 | 30.426 | 1010 | 34E000601 | B | LS1-N2301 | LS1-CB10 | B
CB10 | SCM802 | TX-882A | x110469 | 30.679 | 1010 | 34E000601 | C | LS1-N2301 | LS1-CB10 | C
CB10 | SCM802 | TX-882B | x110452 | 31.413 | 1010 | 34E000601 | D | LS1-N2301 | LS1-CB10 | D
CB10 | SCM802 | TX-883A | x110574 | 30.355 | 1010 | 34E000604 | A | LS1-N2301 | LS1-CB10 | E
CB10 | SCM802 | TX-883B | x110460 | 30.816 | 1010 | 34E000604 | B | LS1-N2301 | LS1-CB10 | F
CB10 | SCM802 | TX-884A | x110463 | 30.689 | 1010 | 34E000604 | C | LS1-N2301 | LS1-CB10 | G
CB10 | SCM802 | TX-884B | x105224 | 34.469 | 1010 | 34E000604 | D | LS1-N2301 | LS1-CB10 | H
CB10 | SCM802 | TX-885A | x110545 | 31.017 | 1010 | 34E000606 | A | LS1-N2301 | LS1-CB10-CB11 | A
CB10 | SCM802 | TX-885B | x110575 | 30.483 | 1010 | 34E000606 | B | LS1-N2301 | LS1-CB10-CB11 | B
CB10 | SCM802 | TX-886A | x109869 | 31.434 | 1010 | 34E000606 | C | LS1-N2301 | LS1-CB10-CB11 | C
CB10 | SCM802 | TX-886B | x110456 | 30.151 | 1010 | 34E000606 | D | LS1-N2301 | LS1-CB10-CB11 | D
CB11 | SCM803 | TX-881A | x110650 | 85.6 | 1050 | 34E000605 | A | LS1-N2301 | LS1-CB10-CB11 | E
CB11 | SCM803 | TX-881B | x114273 | 63.5 | 1050 | 34E000605 | B | LS1-N2301 | LS1-CB10-CB11 | F
CB11 | SCM803 | TX-882A | x112401 | 66.8 | 1050 | 34E000605 | C | LS1-N2301 | LS1-CB10-CB11 | G
CB11 | SCM803 | TX-882B | x114274 | 63.1 | 1050 | 34E000605 | D | LS1-N2301 | LS1-CB10-CB11 | H
CB11 | SCM803 | TX-883A | x110651 | 85.5 | 1050 | 34E000602 | A | LS1-N2301 | LS1-CB11 | A
CB11 | SCM803 | TX-883B | x113702 | 53.5 | 1050 | 34E000602 | B | LS1-N2301 | LS1-CB11 | B
CB11 | SCM803 | TX-884A | x110664 | 83.8 | 1050 | 34E000602 | C | LS1-N2301 | LS1-CB11 | C
CB11 | SCM803 | TX-884B | x110653 | 85.3 | 1050 | 34E000602 | D | LS1-N2301 | LS1-CB11 | D
CB11 | SCM803 | TX-885A | x110665 | 87.8 | 1050 | 34E000603 | A | LS1-N2301 | LS1-CB11 | E
CB11 | SCM803 | TX-885B | x110652 | 93 | 1050 | 34E000603 | B | LS1-N2301 | LS1-CB11 | F
CB11 | SCM803 | TX-886A | x112416 | 76.2 | 1050 | 34E000603 | C | LS1-N2301 | LS1-CB11 | G
CB11 | SCM803 | TX-886B | x110666 | 86 | 1050 | 34E000603 | D | LS1-N2301 | LS1-CB11 | H
CC01 | SCM201 | TX-281A | x113471 | 62.1 | 1050 | 34E000607 | A | LS2-N0208 | LS2-CC01 | A
CC01 | SCM201 | TX-281B | x112355 | 74.9 | 1050 | 34E000607 | B | LS2-N0208 | LS2-CC01 | B
CC01 | SCM201 | TX-282A | x113404 | 62.2 | 1050 | 34E000607 | A | LS2-N0208 | LS2-CC01 | C
CC01 | SCM201 | TX-282B | x113469 | 61.5 | 1050 | 34E000607 | B | LS2-N0208 | LS2-CC01 | D
CC02 | SCM202 | TX-281A | x111759 | 84.5366 | 1050 | 34E000608 | A | LS2-N0208 | LS2-CC01 | E
CC02 | SCM202 | TX-281B | x111768 | 86.4223 | 1050 | 34E000608 | B | LS2-N0208 | LS2-CC01 | F
CC02 | SCM202 | TX-282A | x111677 | 86.092 | 1050 | 34E000608 | A | LS2-N0208 | LS2-CC01 | G
CC02 | SCM202 | TX-282B | x113495 | 55.3134 | 1050 | 34E000608 | B | LS2-N0208 | LS2-CC01 | H
CC03 | SCM203 | TX-281A | x111774 | 84.385 | 1050 | 34E000610 | A | LS2-N0508 | LS2-CC03 | A
CC03 | SCM203 | TX-281B | x111788 | 84.61 | 1050 | 34E000610 | B | LS2-N0508 | LS2-CC03 | B
CC03 | SCM203 | TX-282A | x111675 | 86.1088 | 1050 | 34E000610 | A | LS2-N0508 | LS2-CC03 | C
CC03 | SCM203 | TX-282B | x111745 | 84.2801 | 1050 | 34E000610 | B | LS2-N0508 | LS2-CC03 | D
CC04 | SCM204 | TX-281A | x111785 | 84.26 | 1050 | 34E000609 | A | LS2-N0508 | LS2-CC03 | E
CC04 | SCM204 | TX-281B | x111727 | 83.88 | 1050 | 34E000609 | B | LS2-N0508 | LS2-CC03 | F
CC04 | SCM204 | TX-282A | x111725 | 84.39 | 1050 | 34E000609 | A | LS2-N0508 | LS2-CC03 | G
CC04 | SCM204 | TX-282B | x111728 | 86.03 | 1050 | 34E000609 | B | LS2-N0508 | LS2-CC03 | H
CC05 | SCM205 | TX-281A | x111682 | 83.8122 | 1050 | 34E000612 | A | LS2-N0708 | LS2-CC05 | A
CC05 | SCM205 | TX-281B | x111641 | 85.1023 | 1050 | 34E000612 | B | LS2-N0708 | LS2-CC05 | B
CC05 | SCM205 | TX-282A | x111685 | 86.0957 | 1050 | 34E000612 | A | LS2-N0708 | LS2-CC05 | C
CC05 | SCM205 | TX-282B | x111783 | 84.771 | 1050 | 34E000612 | B | LS2-N0708 | LS2-CC05 | D
CC06 | SCM206 | TX-281A | x111715 | 83.8623 | 1050 | 34E000614 | A | LS2-N0708 | LS2-CC05 | E
CC06 | SCM206 | TX-281B | x111681 | 84.4967 | 1050 | 34E000614 | B | LS2-N0708 | LS2-CC05 | F
CC06 | SCM206 | TX-282A | x111756 | 86.8786 | 1050 | 34E000614 | A | LS2-N0708 | LS2-CC05 | G
CC06 | SCM206 | TX-282B | x110680 | 89.9265 | 1050 | 34E000614 | B | LS2-N0708 | LS2-CC05 | H
CC07 | SCM209 | TX-281A | x113524 | 52.821 | 1050 | 34E000613 | A | LS2-N1008 | LS2-CC07 | A
CC07 | SCM209 | TX-281B | x113521 | 54.0097 | 1050 | 34E000613 | B | LS2-N1008 | LS2-CC07 | B
CC07 | SCM209 | TX-282A | x113518 | 52.8104 | 1050 | 34E000613 | A | LS2-N1008 | LS2-CC07 | C
CC07 | SCM209 | TX-282B | x113516 | 62.45 | 1050 | 34E000613 | B | LS2-N1008 | LS2-CC07 | D
CC08 | SCM208 | TX-281A | x113520 | 55.3248 | 1050 | 34E000611 | A | LS2-N1008 | LS2-CC07 | E
CC08 | SCM208 | TX-281B | x113523 | 54.5217 | 1050 | 34E000611 | B | LS2-N1008 | LS2-CC07 | F
CC08 | SCM208 | TX-282A | x113528 | 52.8702 | 1050 | 34E000611 | A | LS2-N1008 | LS2-CC07 | G
CC08 | SCM208 | TX-282B | x113522 | 57.0874 | 1050 | 34E000611 | B | LS2-N1008 | LS2-CC07 | H
CC09 | SCM210 | TX-281A | x111757 | 52.821 | 1050 | 34E000617 | A | LS2-N1308 | LS2-CC09 | A
CC09 | SCM210 | TX-281B | x111790 | 54.0097 | 1050 | 34E000617 | B | LS2-N1308 | LS2-CC09 | B
CC09 | SCM210 | TX-282A | x111686 | 52.8104 | 1050 | 34E000617 | A | LS2-N1308 | LS2-CC09 | C
CC09 | SCM210 | TX-282B | x111776 | 62.45 | 1050 | 34E000617 | B | LS2-N1308 | LS2-CC09 | D
CC10 | SCM211 | TX-281A | x111748 | 84.2006 | 1050 | 34E000616 | A | LS2-N1308 | LS2-CC09 | E
CC10 | SCM211 | TX-281B | x111760 | 83.8764 | 1050 | 34E000616 | B | LS2-N1308 | LS2-CC09 | F
CC10 | SCM211 | TX-282A | x111670 | 85.2919 | 1050 | 34E000616 | A | LS2-N1308 | LS2-CC09 | G
CC10 | SCM211 | TX-282B | x111786 | 85.2344 | 1050 | 34E000616 | B | LS2-N1308 | LS2-CC09 | H
CC11 | SCM212 | TX-281A | x111743 | 86.22 | 1050 | 34E000615 | A | LS2-N1508 | LS2-CC11 | A
CC11 | SCM212 | TX-281B | x111645 | 83.55 | 1050 | 34E000615 | B | LS2-N1508 | LS2-CC11 | B
CC11 | SCM212 | TX-282A | x111749 | 84.1 | 1050 | 34E000615 | A | LS2-N1508 | LS2-CC11 | C
CC11 | SCM212 | TX-282B | x111767 | 83.87 | 1050 | 34E000615 | B | LS2-N1508 | LS2-CC11 | D
CC12 | SCM207 | TX-281A | x111690 | 85.8432 | 1050 | 34E000618 | A | LS2-N1508 | LS2-CC11 | E
CC12 | SCM207 | TX-281B | x111691 | 85.9542 | 1050 | 34E000618 | B | LS2-N1508 | LS2-CC11 | F
CC12 | SCM207 | TX-282A | x113742 | 55.1145 | 1050 | 34E000618 | A | LS2-N1508 | LS2-CC11 | G
CC12 | SCM207 | TX-282B | x113760 | 56.6671 | 1050 | 34E000618 | B | LS2-N1508 | LS2-CC11 | H

## LR Cables Routing
Info | [7 empty] | Status
Number | Patch Cable name | 4/8 RTD Cable | Rack Number | Connecting Cryo-con | Services Cyromodule # | Longrun Cable #1 | Long Run Cable #2 (8RTD cable only) | Build Date | Test Date | Install Date
1 | 1_N0305_1 | 4 | LS1_N0305 | SCM401 | SCM401 | 34E000568 |  | 2017-09-13 | 2017-09-15 | 2017-09-15
2 | 1_N0305_2 | 4 | LS1_N0305 | SCM401-2 | SCM401 | 34E000571 |  | 2017-09-13 | 2017-09-15 | 2017-09-15
3 | 1_N0305_3 | 4 | LS1_N0305 | SCM402 | SCM402 | 34E000572 |  | 2017-09-13 | 2017-09-15 | 2017-09-15
4 | 1_N0305_4 | 4 | LS1_N0305 | SCM402-2 | SCM402 | 34E000573 |  | 2017-09-13 | 2017-09-15 | 2017-09-15
5 | 1_N0305_5 | 4 | LS1_N0305 | SCM403 | SCM403 | 34E000569 |  | 2017-09-13 | 2017-09-15 | 2017-09-15
6 | 1_N0305_6 | 4 | LS1_N0305 | SCM403-2 | SCM403 | 34E000570 |  | 2017-09-13 | 2017-09-15 | 2017-09-15
7 | 1_N0602_1 | 8 | LS1_N0602 | SCM801 | SCM801 | 34E000574 | 34E000582
8 | 1_N0602_2 | 4 | LS1_N0602 | SCM801/2 | SCM801 | 34E000581
9 | 1_N0602_3 | 8 | LS1_N0602 | SCM802 | SCM802 | 34E000580 | 34E000576
10 | 1_N0602_4 | 4 | LS1_N0602 | SCM801/2 | SCM802 | 34E000579
11 | 1_N0602_5 | 8 | LS1_N0602 | SCM803 | SCM803 | 34E000578 | 34E000577
12 | 1_N0602_6 | 4 | LS1_N0602 | SCM803-2 | SCM803 | 34E000575
13 | 1_N1102_1 | 8 | LS1_N1102 | SCM804 | SCM804 | 34E000584 | 34E000587
14 | 1_N1102_2 | 4 | LS1_N1102 | SCM804/5 | SCM804 | 34E000588
15 | 1_N1102_3 | 8 | LS1_N1102 | SCM805 | SCM805 | 34E000585 | 34E000583
16 | 1_N1102_4 | 4 | LS1_N1102 | SCM804/5 | SCM805 | 34E000586
17 | 1_N1501_1 | 8 | LS1_N1501 | SCM806 | SCM806 | 34E000589 | 34E000591
18 | 1_N1501_2 | 4 | LS1_N1501 | SCM806/7 | SCM806 | 34E000593
19 | 1_N1501_3 | 8 | LS1_N1501 | SCM807 | SCM807 | 34E000594 | 34E000590
20 | 1_N1501_4 | 4 | LS1_N1501 | SCM806/7 | SCM807 | 34E000592
21 | 1_N1701_1 | 8 | LS1_N1701 | SCM808 | SCM808 | 34E000595 | 34E000598
22 | 1_N1701_2 | 4 | LS1_N1701 | SCM808/9 | SCM808 | 34E000597
23 | 1_N1701_3 | 8 | LS1_N1701 | SCM809 | SCM809 | 34E000600 | 34E000599
24 | 1_N1701_4 | 4 | LS1_N1701 | SCM808/9 | SCM809 | 34E000596
25 | 1_N2301_1 | 8 | LS1_N2301 | SCM810 | SCM810 | 34E000601 | 34E000604
26 | 1_N2301_2 | 4 | LS1_N2301 | SCM810/11 | SCM810 | 34E000606
27 | 1_N2301_3 | 8 | LS1_N2301 | SCM811 | SCM811 | 34E000605 | 34E000602
28 | 1_N2301_4 | 4 | LS1_N2301 | SCM810/11 | SCM811 | 34E000603
29 | 2_N0208_1 | 4 | LS2_N0208 | SCM201 | SCM201 | 34E000607
30 | 2_N0208_2 | 4 | LS2_N0208 | SCM202 | SCM202 | 34E000608
31 | 2_N0508_1 | 4 | LS2_N0508 | SCM203 | SCM203 | 34E000610
32 | 2_N0508_2 | 4 | LS2_N0508 | SCM204 | SCM204 | 34E000609
33 | 2_N0708_1 | 4 | LS2_N0708 | SCM205 | SCM205 | 34E000612
34 | 2_N0708_2 | 4 | LS2_N0708 | SCM206 | SCM206 | 34E000614
35 | 2_N1008_1 | 4 | LS2_N1008 | SCM207 | SCM207 | 34E000613
36 | 2_N1008_2 | 4 | LS2_N1008 | SCM208 | SCM208 | 34E000611
37 | 2_N1308_1 | 4 | LS2_N1308 | SCM209 | SCM209 | 34E000617
38 | 2_N1308_2 | 4 | LS2_N1308 | SCM210 | SCM210 | 34E000616
39 | 2_N1508_1 | 4 | LS2_N1508 | SCM211 | SCM211 | 34E000615
40 | 2_N1508_2 | 4 | LS2_N1508 | SCM212 | SCM212 | 34E000618

## Testing Worksheets
Diagnostic RTD Installation Data
Beam Line Section | CyroModule Type | CyroModule Number | RTD Location Tag | RTD Install Location | RTD Serial Number | Feedthough Port | Resistance @ room temp from datasheet | Resistance Tested after install at Port | 2 wire resistance test | 4 wire resistance test | Notes
LS2 | Beta .29 HWR | #10 | TX-281A | Before Solenoid 1 | x111757 | A5 | 84.2006 | 83.9
LS2 | Beta .29 HWR | #10 | TX-281B | Before Solenoid 1 | x111790 | A5 | 83.8764 | 83.6
LS2 | Beta .29 HWR | #10 | TX-282A | After Solenoid 1 | x111686 | A5 | 85.2919 | 85
LS2 | Beta .29 HWR | #10 | TX-282B | After Solenoid 1 | x111776 | A5 | 85.2344 | 84.9
LS2 | Beta .29 HWR | #12 | TX-281A | Before Solenoid 1 | x111743 | A5 | 86.2905 | 86.2
LS2 | Beta .29 HWR | #12 | TX-281B | Before Solenoid 1 | x111645 | A5 | 84.58 | 84.4
LS2 | Beta .29 HWR | #12 | TX-282A | After Solenoid 1 | x111749 | A5 | 82.8966 | 82.9
LS2 | Beta .29 HWR | #12 | TX-282B | After Solenoid 1 | x111767 | A5 | 85.0437 | 84.9
LS1 | Beta .085 QWR | #12 | TX-881A | Before Solenoid 1 | x113738 | A5 | 54.78 | 54.9 Ω, OK
LS1 | Beta .085 QWR | #12 | TX-881B | Before Solenoid 1 | x113710 | A5 | 57.6569 | 57.7 Ω, OK
LS1 | Beta .085 QWR | #12 | TX-882A | After Solenoid 1 | x113816 | A5 | 59.1895 | 59.1 Ω, OK
LS1 | Beta .085 QWR | #12 | TX-882B | After Solenoid 1 | x113754 | A5 | 59.5311 | 59.5 Ω, OK
LS1 | Beta .085 QWR | #12 | TX-883A | Before Solenoid 2 | x113746 | A5 | 55.1295 | 55.1 Ω, OK
LS1 | Beta .085 QWR | #12 | TX-883B | Before Solenoid 2 | x113504 | A5 | 62.1356 | 62.1 Ω, OK
LS1 | Beta .085 QWR | #12 | TX-884A | After Solenoid 2 | x113820 | A5 | 58.9231 | 58.7 Ω, OK
LS1 | Beta .085 QWR | #12 | TX-884B | After Solenoid 2 | x113503 | A5 | 61.7911 | 61.5 Ω, OK
LS1 | Beta .085 QWR | #12 | TX-885A | Before Solenoid 3 | x113751 | A6 | 57.5281 | 57.5 Ω, OK
LS1 | Beta .085 QWR | #12 | TX-885B | Before Solenoid 3 | x113747 | A6 | 57.3114 | 57.3 Ω, OK
LS1 | Beta .085 QWR | #12 | TX-886A | After Solenoid 3 | x113723 | A6 | 55.4193 | 55.3 Ω, OK
LS1 | Beta .085 QWR | #12 | TX-886B | After Solenoid 3 | x113800 | A6 | 56.2635 | 56.2 Ω, OK

## Quick Cabling Guide
Pin | Cable | Color
1 | A | GROUND of 'white cables' (pins 2-9)
2 | A | White/Brown
3 | A | Brown/White
4 | A | White/Green
5 | A | Green/White
6 | A | White/Orange
7 | A | Orange/White
8 | A | White/Blue
9 | A | Blue/White
10 | A | GROUND of 'red cables' (pins 11-18)
11 | A | Red/Brown
12 | A | Brown/Red
13 | A | Red/Green
14 | A | Green/Red
15 | A | Red/Orange
16 | A | Orange/Red
17 | A | Red/Blue
18 | A | Blue/Red
19 | B | GROUND of 'white cables' (pins 20-27)
20 | B | White/Brown
21 | B | Brown/White
22 | B | White/Green
23 | B | Green/White
24 | B | White/Orange
25 | B | Orange/White
26 | B | White/Blue
27 | B | Blue/White
28 | B | GROUND of 'red cables' (pins 29-36)
29 | B | Red/Brown
30 | B | Brown/Red
31 | B | Red/Green
32 | B | Green/Red
33 | B | Red/Orange
34 | B | Orange/Red
35 | B | Red/Blue
36 | B | Blue/Red
37

## Sheet2
LS1:CB01: | 1 | LS1: | CB01: | FTHS:TX881A | LS1:CB01:FTHS:TX881A
LS1:CB01:FTS:TX881B | 2 | LS1: | CB01: | FTHS:TX881B | LS1:CB01:FTHS:TX881B
LS1:CB01:FTS:TX882A | 3 | LS1: | CB01: | FTHS:TX882A | LS1:CB01:FTHS:TX882A
LS1:CB01:FTS:TX882B | 4 | LS1: | CB01: | FTHS:TX882B | LS1:CB01:FTHS:TX882B
LS1:CB01:FTS:TX883A | 5 | LS1: | CB01: | FTHS:TX883A | LS1:CB01:FTHS:TX883A
LS1:CB01:FTS:TX883B | 6 | LS1: | CB01: | FTHS:TX883B | LS1:CB01:FTHS:TX883B
LS1:CB01:FTS:TX884A | 7 | LS1: | CB01: | FTHS:TX884A | LS1:CB01:FTHS:TX884A
LS1:CB01:FTS:TX884B | 8 | LS1: | CB01: | FTHS:TX884B | LS1:CB01:FTHS:TX884B
LS1:CB01:FTS:TX885A | 9 | LS1: | CB01: | FTHS:TX885A | LS1:CB01:FTHS:TX885A
LS1:CB01:FTS:TX885B | 10 | LS1: | CB01: | FTHS:TX885B | LS1:CB01:FTHS:TX885B
LS1:CB01:FTS:TX886A | 11 | LS1: | CB01: | FTHS:TX886A | LS1:CB01:FTHS:TX886A
LS1:CB01:FTS:TX886B | 12 | LS1: | CB01: | FTHS:TX886A | LS1:CB01:FTHS:TX886A
LS1:CB02:FTS:TX881A | 1 | LS1: | CB02: | FTHS:TX881A | LS1:CB02:FTHS:TX881A
LS1:CB02:FTS:TX881B | 2 | LS1: | CB02: | FTHS:TX881B | LS1:CB02:FTHS:TX881B
LS1:CB02:FTS:TX882A | 3 | LS1: | CB02: | FTHS:TX882A | LS1:CB02:FTHS:TX882A
LS1:CB02:FTS:TX882B | 4 | LS1: | CB02: | FTHS:TX882B | LS1:CB02:FTHS:TX882B
LS1:CB03:FTS:TX883A | 5 | LS1: | CB02: | FTHS:TX883A | LS1:CB02:FTHS:TX883A
LS1:CB03:FTS:TX883B | 6 | LS1: | CB02: | FTHS:TX883B | LS1:CB02:FTHS:TX883B
LS1:CB03:FTS:TX884A | 7 | LS1: | CB02: | FTHS:TX884A | LS1:CB02:FTHS:TX884A
LS1:CB03:FTS:TX884B | 8 | LS1: | CB02: | FTHS:TX884B | LS1:CB02:FTHS:TX884B
LS1:CB03:FTS:TX885A | 9 | LS1: | CB02: | FTHS:TX885A | LS1:CB02:FTHS:TX885A
LS1:CB03:FTS:TX885B | 10 | LS1: | CB02: | FTHS:TX885B | LS1:CB02:FTHS:TX885B
LS1:CB03:FTS:TX886A | 11 | LS1: | CB02: | FTHS:TX886A | LS1:CB02:FTHS:TX886A
LS1:CB03:FTS:TX886B | 12 | LS1: | CB02: | FTHS:TX886A | LS1:CB02:FTHS:TX886A
LS1:CB04:FTS:TX881A | 1 | LS1: | CB03: | FTHS:TX881A | LS1:CB03:FTHS:TX881A
LS1:CB04:FTS:TX881B | 2 | LS1: | CB03: | FTHS:TX881B | LS1:CB03:FTHS:TX881B
LS1:CB04:FTS:TX882A | 3 | LS1: | CB03: | FTHS:TX882A | LS1:CB03:FTHS:TX882A
LS1:CB04:FTS:TX882B | 4 | LS1: | CB03: | FTHS:TX882B | LS1:CB03:FTHS:TX882B
LS1:CB04:FTS:TX883A | 5 | LS1: | CB03: | FTHS:TX883A | LS1:CB03:FTHS:TX883A
LS1:CB04:FTS:TX883B | 6 | LS1: | CB03: | FTHS:TX883B | LS1:CB03:FTHS:TX883B
LS1:CB04:FTS:TX884A | 7 | LS1: | CB03: | FTHS:TX884A | LS1:CB03:FTHS:TX884A
LS1:CB04:FTS:TX884B | 8 | LS1: | CB03: | FTHS:TX884B | LS1:CB03:FTHS:TX884B
LS1:CB04:FTS:TX885A | 9 | LS1: | CB03: | FTHS:TX885A | LS1:CB03:FTHS:TX885A
LS1:CB04:FTS:TX885B | 10 | LS1: | CB03: | FTHS:TX885B | LS1:CB03:FTHS:TX885B
LS1:CB04:FTS:TX886A | 11 | LS1: | CB03: | FTHS:TX886A | LS1:CB03:FTHS:TX886A
LS1:CB04:FTS:TX886B | 12 | LS1: | CB03: | FTHS:TX886A | LS1:CB03:FTHS:TX886A
LS1:CB04:FTS:TX881A | 1 | LS1: | CB04: | FTHS:TX881A | LS1:CB04:FTHS:TX881A
LS1:CB04:FTS:TX881B | 2 | LS1: | CB04: | FTHS:TX881B | LS1:CB04:FTHS:TX881B
LS1:CB04:FTS:TX882A | 3 | LS1: | CB04: | FTHS:TX882A | LS1:CB04:FTHS:TX882A
LS1:CB04:FTS:TX882B | 4 | LS1: | CB04: | FTHS:TX882B | LS1:CB04:FTHS:TX882B
LS1:CB04:FTS:TX883A | 5 | LS1: | CB04: | FTHS:TX883A | LS1:CB04:FTHS:TX883A
LS1:CB04:FTS:TX883B | 6 | LS1: | CB04: | FTHS:TX883B | LS1:CB04:FTHS:TX883B
LS1:CB04:FTS:TX884A | 7 | LS1: | CB04: | FTHS:TX884A | LS1:CB04:FTHS:TX884A
LS1:CB04:FTS:TX884B | 8 | LS1: | CB04: | FTHS:TX884B | LS1:CB04:FTHS:TX884B
LS1:CB04:FTS:TX885A | 9 | LS1: | CB04: | FTHS:TX885A | LS1:CB04:FTHS:TX885A
LS1:CB04:FTS:TX885B | 10 | LS1: | CB04: | FTHS:TX885B | LS1:CB04:FTHS:TX885B
LS1:CB04:FTS:TX886A | 11 | LS1: | CB04: | FTHS:TX886A | LS1:CB04:FTHS:TX886A
LS1:CB04:FTS:TX886B | 12 | LS1: | CB04: | FTHS:TX886A | LS1:CB04:FTHS:TX886A
LS1:CB05:FTS:TX881A | 1 | LS1: | CB05: | FTHS:TX881A | LS1:CB05:FTHS:TX881A
LS1:CB05:FTS:TX881B | 2 | LS1: | CB05: | FTHS:TX881B | LS1:CB05:FTHS:TX881B
LS1:CB05:FTS:TX882A | 3 | LS1: | CB05: | FTHS:TX882A | LS1:CB05:FTHS:TX882A
LS1:CB05:FTS:TX882B | 4 | LS1: | CB05: | FTHS:TX882B | LS1:CB05:FTHS:TX882B
LS1:CB05:FTS:TX883A | 5 | LS1: | CB05: | FTHS:TX883A | LS1:CB05:FTHS:TX883A
LS1:CB05:FTS:TX883B | 6 | LS1: | CB05: | FTHS:TX883B | LS1:CB05:FTHS:TX883B
LS1:CB05:FTS:TX884A | 7 | LS1: | CB05: | FTHS:TX884A | LS1:CB05:FTHS:TX884A
LS1:CB05:FTS:TX884B | 8 | LS1: | CB05: | FTHS:TX884B | LS1:CB05:FTHS:TX884B
LS1:CB05:FTS:TX885A | 9 | LS1: | CB05: | FTHS:TX885A | LS1:CB05:FTHS:TX885A
LS1:CB05:FTS:TX885B | 10 | LS1: | CB05: | FTHS:TX885B | LS1:CB05:FTHS:TX885B
LS1:CB05:FTS:TX886A | 11 | LS1: | CB05: | FTHS:TX886A | LS1:CB05:FTHS:TX886A
LS1:CB05:FTS:TX886B | 12 | LS1: | CB05: | FTHS:TX886A | LS1:CB05:FTHS:TX886A
LS1:CB06:FTS:TX881A | 1 | LS1: | CB06: | FTHS:TX881A | LS1:CB06:FTHS:TX881A
LS1:CB06:FTS:TX881B | 2 | LS1: | CB06: | FTHS:TX881B | LS1:CB06:FTHS:TX881B
LS1:CB06:FTS:TX882A | 3 | LS1: | CB06: | FTHS:TX882A | LS1:CB06:FTHS:TX882A
LS1:CB06:FTS:TX882B | 4 | LS1: | CB06: | FTHS:TX882B | LS1:CB06:FTHS:TX882B
LS1:CB06:FTS:TX883A | 5 | LS1: | CB06: | FTHS:TX883A | LS1:CB06:FTHS:TX883A
LS1:CB06:FTS:TX883B | 6 | LS1: | CB06: | FTHS:TX883B | LS1:CB06:FTHS:TX883B
LS1:CB06:FTS:TX884A | 7 | LS1: | CB06: | FTHS:TX884A | LS1:CB06:FTHS:TX884A
LS1:CB06:FTS:TX884B | 8 | LS1: | CB06: | FTHS:TX884B | LS1:CB06:FTHS:TX884B
LS1:CB06:FTS:TX885A | 9 | LS1: | CB06: | FTHS:TX885A | LS1:CB06:FTHS:TX885A
LS1:CB06:FTS:TX885B | 10 | LS1: | CB06: | FTHS:TX885B | LS1:CB06:FTHS:TX885B
LS1:CB06:FTS:TX886A | 11 | LS1: | CB06: | FTHS:TX886A | LS1:CB06:FTHS:TX886A
LS1:CB06:FTS:TX886B | 12 | LS1: | CB06: | FTHS:TX886A | LS1:CB06:FTHS:TX886A
LS1:CB07:FTS:TX881A | 1 | LS1: | CB07: | FTHS:TX881A | LS1:CB07:FTHS:TX881A
LS1:CB07:FTS:TX881B | 2 | LS1: | CB07: | FTHS:TX881B | LS1:CB07:FTHS:TX881B
LS1:CB07:FTS:TX882A | 3 | LS1: | CB07: | FTHS:TX882A | LS1:CB07:FTHS:TX882A
LS1:CB07:FTS:TX882B | 4 | LS1: | CB07: | FTHS:TX882B | LS1:CB07:FTHS:TX882B
LS1:CB07:FTS:TX883A | 5 | LS1: | CB07: | FTHS:TX883A | LS1:CB07:FTHS:TX883A
LS1:CB07:FTS:TX883B | 6 | LS1: | CB07: | FTHS:TX883B | LS1:CB07:FTHS:TX883B
LS1:CB07:FTS:TX884A | 7 | LS1: | CB07: | FTHS:TX884A | LS1:CB07:FTHS:TX884A
LS1:CB07:FTS:TX884B | 8 | LS1: | CB07: | FTHS:TX884B | LS1:CB07:FTHS:TX884B
LS1:CB07:FTS:TX885A | 9 | LS1: | CB07: | FTHS:TX885A | LS1:CB07:FTHS:TX885A
LS1:CB07:FTS:TX885B | 10 | LS1: | CB07: | FTHS:TX885B | LS1:CB07:FTHS:TX885B
LS1:CB07:FTS:TX886A | 11 | LS1: | CB07: | FTHS:TX886A | LS1:CB07:FTHS:TX886A
LS1:CB07:FTS:TX886B | 12 | LS1: | CB07: | FTHS:TX886A | LS1:CB07:FTHS:TX886A
LS1:CB08:FTS:TX881A | 1 | LS1: | CB08: | FTHS:TX881A | LS1:CB08:FTHS:TX881A
LS1:CB08:FTS:TX881B | 2 | LS1: | CB08: | FTHS:TX881B | LS1:CB08:FTHS:TX881B
LS1:CB08:FTS:TX882A | 3 | LS1: | CB08: | FTHS:TX882A | LS1:CB08:FTHS:TX882A
LS1:CB08:FTS:TX882B | 4 | LS1: | CB08: | FTHS:TX882B | LS1:CB08:FTHS:TX882B
LS1:CB08:FTS:TX883A | 5 | LS1: | CB08: | FTHS:TX883A | LS1:CB08:FTHS:TX883A
LS1:CB08:FTS:TX883B | 6 | LS1: | CB08: | FTHS:TX883B | LS1:CB08:FTHS:TX883B
LS1:CB08:FTS:TX884A | 7 | LS1: | CB08: | FTHS:TX884A | LS1:CB08:FTHS:TX884A
LS1:CB08:FTS:TX884B | 8 | LS1: | CB08: | FTHS:TX884B | LS1:CB08:FTHS:TX884B
LS1:CB08:FTS:TX885A | 9 | LS1: | CB08: | FTHS:TX885A | LS1:CB08:FTHS:TX885A
LS1:CB08:FTS:TX885B | 10 | LS1: | CB08: | FTHS:TX885B | LS1:CB08:FTHS:TX885B
LS1:CB08:FTS:TX886A | 11 | LS1: | CB08: | FTHS:TX886A | LS1:CB08:FTHS:TX886A
LS1:CB08:FTS:TX886B | 12 | LS1: | CB08: | FTHS:TX886A | LS1:CB08:FTHS:TX886A
LS1:CB09:FTS:TX881A | 1 | LS1: | CB09: | FTHS:TX881A | LS1:CB09:FTHS:TX881A
LS1:CB09:FTS:TX881B | 2 | LS1: | CB09: | FTHS:TX881B | LS1:CB09:FTHS:TX881B
LS1:CB09:FTS:TX882A | 3 | LS1: | CB09: | FTHS:TX882A | LS1:CB09:FTHS:TX882A
LS1:CB09:FTS:TX882B | 4 | LS1: | CB09: | FTHS:TX882B | LS1:CB09:FTHS:TX882B
LS1:CB09:FTS:TX883A | 5 | LS1: | CB09: | FTHS:TX883A | LS1:CB09:FTHS:TX883A
LS1:CB09:FTS:TX883B | 6 | LS1: | CB09: | FTHS:TX883B | LS1:CB09:FTHS:TX883B
LS1:CB09:FTS:TX884A | 7 | LS1: | CB09: | FTHS:TX884A | LS1:CB09:FTHS:TX884A
LS1:CB09:FTS:TX884B | 8 | LS1: | CB09: | FTHS:TX884B | LS1:CB09:FTHS:TX884B
LS1:CB09:FTS:TX885A | 9 | LS1: | CB09: | FTHS:TX885A | LS1:CB09:FTHS:TX885A
LS1:CB09:FTS:TX885B | 10 | LS1: | CB09: | FTHS:TX885B | LS1:CB09:FTHS:TX885B
LS1:CB09:FTS:TX886A | 11 | LS1: | CB09: | FTHS:TX886A | LS1:CB09:FTHS:TX886A
LS1:CB09:FTS:TX886B | 12 | LS1: | CB09: | FTHS:TX886A | LS1:CB09:FTHS:TX886A
LS1:CB10:FTS:TX881A | 1 | LS1: | CB09: | FTHS:TX881A | LS1:CB09:FTHS:TX881A
LS1:CB10:FTS:TX881B | 2 | LS1: | CB09: | FTHS:TX881B | LS1:CB09:FTHS:TX881B
LS1:CB10:FTS:TX882A | 3 | LS1: | CB09: | FTHS:TX882A | LS1:CB09:FTHS:TX882A
LS1:CB10:FTS:TX882B | 4 | LS1: | CB09: | FTHS:TX882B | LS1:CB09:FTHS:TX882B
LS1:CB10:FTS:TX883A | 5 | LS1: | CB09: | FTHS:TX883A | LS1:CB09:FTHS:TX883A
LS1:CB10:FTS:TX883B | 6 | LS1: | CB09: | FTHS:TX883B | LS1:CB09:FTHS:TX883B
LS1:CB10:FTS:TX884A | 7 | LS1: | CB09: | FTHS:TX884A | LS1:CB09:FTHS:TX884A
LS1:CB10:FTS:TX884B | 8 | LS1: | CB09: | FTHS:TX884B | LS1:CB09:FTHS:TX884B
LS1:CB10:FTS:TX885A | 9 | LS1: | CB09: | FTHS:TX885A | LS1:CB09:FTHS:TX885A
LS1:CB10:FTS:TX885B | 10 | LS1: | CB09: | FTHS:TX885B | LS1:CB09:FTHS:TX885B
LS1:CB10:FTS:TX886A | 11 | LS1: | CB09: | FTHS:TX886A | LS1:CB09:FTHS:TX886A
LS1:CB10:FTS:TX886B | 12 | LS1: | CB09: | FTHS:TX886A | LS1:CB09:FTHS:TX886A
LS1:CB11:FTS:TX881A | 1 | LS1: | CB10: | FTHS:TX881A | LS1:CB10:FTHS:TX881A
LS1:CB11:FTS:TX881B | 2 | LS1: | CB10: | FTHS:TX881B | LS1:CB10:FTHS:TX881B
LS1:CB11:FTS:TX882A | 3 | LS1: | CB10: | FTHS:TX882A | LS1:CB10:FTHS:TX882A
LS1:CB11:FTS:TX882B | 4 | LS1: | CB10: | FTHS:TX882B | LS1:CB10:FTHS:TX882B
LS1:CB11:FTS:TX883A | 5 | LS1: | CB10: | FTHS:TX883A | LS1:CB10:FTHS:TX883A
LS1:CB11:FTS:TX883B | 6 | LS1: | CB10: | FTHS:TX883B | LS1:CB10:FTHS:TX883B
LS1:CB11:FTS:TX884A | 7 | LS1: | CB10: | FTHS:TX884A | LS1:CB10:FTHS:TX884A
LS1:CB11:FTS:TX884B | 8 | LS1: | CB10: | FTHS:TX884B | LS1:CB10:FTHS:TX884B
LS1:CB11:FTS:TX885A | 9 | LS1: | CB10: | FTHS:TX885A | LS1:CB10:FTHS:TX885A
LS1:CB11:FTS:TX885B | 10 | LS1: | CB10: | FTHS:TX885B | LS1:CB10:FTHS:TX885B
LS1:CB11:FTS:TX886A | 11 | LS1: | CB10: | FTHS:TX886A | LS1:CB10:FTHS:TX886A
LS1:CB11:FTS:TX886B | 12 | LS1: | CB10: | FTHS:TX886A | LS1:CB10:FTHS:TX886A
LS1:CC01:FTS:TX281A | 1 | LS1: | CB11: | FTHS:TX881A | LS1:CB11:FTHS:TX881A
LS1:CC01:FTS:TX281B | 2 | LS1: | CB11: | FTHS:TX881B | LS1:CB11:FTHS:TX881B
LS1:CC01:FTS:TX282A | 3 | LS1: | CB11: | FTHS:TX882A | LS1:CB11:FTHS:TX882A
LS1:CC01:FTS:TX282B | 4 | LS1: | CB11: | FTHS:TX882B | LS1:CB11:FTHS:TX882B
LS1:CC02:FTS:TX281A | 5 | LS1: | CB11: | FTHS:TX883A | LS1:CB11:FTHS:TX883A
LS1:CC02:FTS:TX281B | 6 | LS1: | CB11: | FTHS:TX883B | LS1:CB11:FTHS:TX883B
LS1:CC02:FTS:TX282A | 7 | LS1: | CB11: | FTHS:TX884A | LS1:CB11:FTHS:TX884A
LS1:CC02:FTS:TX282B | 8 | LS1: | CB11: | FTHS:TX884B | LS1:CB11:FTHS:TX884B
LS1:CC03:FTS:TX281A | 9 | LS1: | CB11: | FTHS:TX885A | LS1:CB11:FTHS:TX885A
LS1:CC03:FTS:TX281B | 10 | LS1: | CB11: | FTHS:TX885B | LS1:CB11:FTHS:TX885B
LS1:CC03:FTS:TX282A | 11 | LS1: | CB11: | FTHS:TX886A | LS1:CB11:FTHS:TX886A
LS1:CC03:FTS:TX282B | 12 | LS1: | CB11: | FTHS:TX886A | LS1:CB11:FTHS:TX886A
LS1:CC04:FTS:TX281A
LS1:CC04:FTS:TX281B
LS1:CC04:FTS:TX282A
LS1:CC04:FTS:TX282B
LS1:CC05:FTS:TX281A
LS1:CC05:FTS:TX281B
LS1:CC05:FTS:TX282A
LS1:CC05:FTS:TX282B
LS1:CC06:FTS:TX281A
LS1:CC06:FTS:TX281B
LS1:CC06:FTS:TX282A
LS1:CC06:FTS:TX282B
LS1:CC07:FTS:TX281A
LS1:CC07:FTS:TX281B
LS1:CC07:FTS:TX282A
LS1:CC07:FTS:TX282B
LS1:CC08:FTS:TX281A
LS1:CC08:FTS:TX281B
LS1:CC08:FTS:TX282A
LS1:CC08:FTS:TX282B
LS1:CC09:FTS:TX281A
LS1:CC09:FTS:TX281B
LS1:CC09:FTS:TX282A
LS1:CC09:FTS:TX282B
LS1:CC10:FTS:TX281A
LS1:CC10:FTS:TX281B
LS1:CC10:FTS:TX282A
LS1:CC10:FTS:TX282B
LS1:CC11:FTS:TX281A
LS1:CC11:FTS:TX281B
LS1:CC11:FTS:TX282A
LS1:CC11:FTS:TX282B
LS1:CC12:FTS:TX281A
LS1:CC12:FTS:TX281B
LS1:CC12:FTS:TX282A
LS1:CC12:FTS:TX282B